## Notes
Notes on this run:
DXE samples from 2025 mesocosms.
Dilutions from this set were run with LTREB batch 6.
PO4 check standards coming out slightly low compared to previous runs, likely because several low standards were excluded from the P curve.
Will need to exclude samples with algae chunks from this run- algae not fully homogenized, and should have been included in the sample.
If a sample was run in the 20260320 set, use the data from that rerun and not from this file!

## QAQC
Detection Date | Detection Time | Run File Name | Sample ID | Cup Number | Sample Type | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION FACTOR | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- dilution corrected | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE NO3-N: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- dilution corrected | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ | original order
STANDARDS:
2026-02-27 | 12:07 | OM_2-27-2026_12-05-56PM.OMN | dye | S8 | Unknown | 1 | 1 | TP | ug P/L | 0.133 | 1.66 | 30.2 | TN | ug N/L | -0.00663 | -0.0863 | -113 |  | 1 | 1 |  | 30.2 | 30.2 | 30.2 | [6 empty] | 1 |  | -113 | -113 | -113 | [6 empty] | 1
2026-02-27 | 12:09 | OM_2-27-2026_12-05-56PM.OMN | dye | S8 | Unknown | 1 | 1 | TP | ug P/L | 0.133 | 1.66 | 30.4 | TN | ug N/L | 0.00637 | 0.155 | -87.1 |  | 1 | 1 |  | 30.4 | 30.4 | 30.4 | [6 empty] | 1 |  | -87.1 | -87.1 | -87.1 | [6 empty] | 2
2026-02-27 | 12:10 | OM_2-27-2026_12-05-56PM.OMN | NO2 | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.135 | 1.67 | 30.6 | TN | ug N/L | 1.98 | 26 | 2580 | 95.3846 | 1 | 1 |  | 30.6 | 30.6 | 30.6 | [6 empty] | 1 |  | 2580 | 2580 | 2580 | [6 empty] | 3
2026-02-27 | 12:11 | OM_2-27-2026_12-05-56PM.OMN | NO3 | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.134 | 1.58 | 27.7 | TN | ug N/L | 1.95 | 24.8 | 2460 |  | 1 | 1 |  | 27.7 | 27.7 | 27.7 | [6 empty] | 1 |  | 2460 | 2460 | 2460 | [6 empty] | 4
2026-02-27 | 12:13 | OM_2-27-2026_12-05-56PM.OMN | NO2 | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.135 | 1.66 | 30.4 | TN | ug N/L | 1.88 | 24.6 | 2450 | 101.626 | 1 | 1 |  | 30.4 | 30.4 | 30.4 | [6 empty] | 1 |  | 2450 | 2450 | 2450 | [6 empty] | 5
2026-02-27 | 12:14 | OM_2-27-2026_12-05-56PM.OMN | NO3 | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.139 | 1.73 | 32.4 | TN | ug N/L | 1.95 | 25 | 2480 |  | 1 | 1 |  | 32.4 | 32.4 | 32.4 | [6 empty] | 1 |  | 2480 | 2480 | 2480 | [6 empty] | 6
2026-02-27 | 12:15 | OM_2-27-2026_12-05-56PM.OMN | water | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.138 | 1.69 | 31.2 | TN | ug N/L | 0.00705 | 0.16 | -86.6 |  | 1 | 1 |  | 31.2 | 31.2 | 31.2 | [6 empty] | 1 |  | -86.6 | -86.6 | -86.6 | [6 empty] | 7
2026-02-27 | 12:16 | OM_2-27-2026_12-05-56PM.OMN | water | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.134 | 1.63 | 29.5 | TN | ug N/L | 0.0153 | -0.0446 | -108 |  | 1 | 1 |  | 29.5 | 29.5 | 29.5 | [6 empty] | 1 |  | -108 | -108 | -108 | [6 empty] | 8
2026-02-27 | 12:18 | OM_2-27-2026_12-05-56PM.OMN | bisulfate carrier | S4 | Unknown | 1 | 1 | TP | ug P/L | 0.00979 | 0.114 | -16.4 | TN | ug N/L | -0.0156 | -0.17 | -121 |  | 1 | 1 |  | -16.4 | -16.4 | -16.4 | [6 empty] | 1 |  | -121 | -121 | -121 | [6 empty] | 9
2026-02-27 | 12:19 | OM_2-27-2026_12-05-56PM.OMN | bisulfate carrier | S4 | Unknown | 1 | 1 | TP | ug P/L | 0.0102 | 0.113 | -16.4 | TN | ug N/L | 0.0184 | 0.378 | -63.6 |  | 1 | 1 |  | -16.4 | -16.4 | -16.4 | [6 empty] | 1 |  | -63.6 | -63.6 | -63.6 | [6 empty] | 10
2026-02-27 | 12:20 | OM_2-27-2026_12-05-56PM.OMN | cal std 150 | 1 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.288 | 5.46 | 150 | TN | ug N/L | 1.89 | 24.1 | 1500 |  | 1 | 1 |  | 150 | 150 | 150 | [6 empty] | 1 |  | 1500 | 1500 | 1500 | [6 empty] | 11
2026-02-27 | 12:22 | OM_2-27-2026_12-05-56PM.OMN | cal std 150 | 1 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.289 | 5.5 | 150 | TN | ug N/L | 1.95 | 24.8 | 1500 |  | 1 | 1 |  | 150 | 150 | 150 | [6 empty] | 1 |  | 1500 | 1500 | 1500 | [6 empty] | 12
2026-02-27 | 12:23 | OM_2-27-2026_12-05-56PM.OMN | cal std 100 | 3 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.191 | 3.67 | 100 | TN | ug N/L | 1.36 | 17.4 | 1000 |  | 1 | 1 |  | 100 | 100 | 100 | [6 empty] | 1 |  | 1000 | 1000 | 1000 | [6 empty] | 13
2026-02-27 | 12:24 | OM_2-27-2026_12-05-56PM.OMN | cal std 100 | 3 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.192 | 3.73 | 100 | TN | ug N/L | 1.31 | 16.8 | 1000 |  | 1 | 1 |  | 100 | 100 | 100 | [6 empty] | 1 |  | 1000 | 1000 | 1000 | [6 empty] | 14
2026-02-27 | 12:25 | OM_2-27-2026_12-05-56PM.OMN | cal std 50 | 5 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0998 | 2.06 | 50 | TN | ug N/L | 0.717 | 9.25 | 500 |  | 1 | 1 |  | 50 | 50 | 50 | [6 empty] | 1 |  | 500 | 500 | 500 | [6 empty] | 15
2026-02-27 | 12:27 | OM_2-27-2026_12-05-56PM.OMN | cal std 50 | 5 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0992 | 2.24 | 50 | TN | ug N/L | 0.696 | 8.98 | 500 |  | 1 | 3 | excluded from curve | 50 | 50 | 50 | [6 empty] | 1 |  | 500 | 500 | 500 | [6 empty] | 16
2026-02-27 | 12:28 | OM_2-27-2026_12-05-56PM.OMN | cal std 25 | 7 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0505 | 1.16 | 25 | TN | ug N/L | 0.402 | 5.26 | 250 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 250 | 250 | 250 | [6 empty] | 17
2026-02-27 | 12:29 | OM_2-27-2026_12-05-56PM.OMN | cal std 25 | 7 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0501 | 1.13 | 25 | TN | ug N/L | 0.387 | 5.08 | 250 |  | 1 | 1 |  | 25 | 25 | 25 | [6 empty] | 1 |  | 250 | 250 | 250 | [6 empty] | 18
2026-02-27 | 12:30 | OM_2-27-2026_12-05-56PM.OMN | cal std 10 | 9 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0231 | 0.611 | 10 | TN | ug N/L | 0.19 | 2.61 | 100 |  | 1 | 1 |  | 10 | 10 | 10 | [6 empty] | 1 |  | 100 | 100 | 100 | [6 empty] | 19
2026-02-27 | 12:32 | OM_2-27-2026_12-05-56PM.OMN | cal std 10 | 9 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0333 | 0.828 | 10 | TN | ug N/L | 0.196 | 2.69 | 100 |  | 1 | 3 | excluded from curve | 10 | 10 | 10 | [6 empty] | 1 |  | 100 | 100 | 100 | [6 empty] | 20
2026-02-27 | 12:33 | OM_2-27-2026_12-05-56PM.OMN | cal std 5 | 11 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0267 | 0.64 | 5 | TN | ug N/L | 0.154 | 2.17 | 50 |  | 1 | 3 | excluded from curve | 5 | 5 | 5 | [6 empty] | 1 |  | 50 | 50 | 50 | [6 empty] | 21
2026-02-27 | 12:34 | OM_2-27-2026_12-05-56PM.OMN | cal std 5 | 11 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0188 | 0.495 | 5 | TN | ug N/L | 0.131 | 1.87 | 50 |  | 1 | 1 |  | 5 | 5 | 5 | [6 empty] | 1 |  | 50 | 50 | 50 | [6 empty] | 22
2026-02-27 | 12:36 | OM_2-27-2026_12-05-56PM.OMN | cal std 0 | 13 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0113 | 0.274 | 0 | TN | ug N/L | 0.0691 | 1.07 | 0 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | 0 | 0 | 0 | [6 empty] | 23
2026-02-27 | 12:37 | OM_2-27-2026_12-05-56PM.OMN | cal std 0 | 13 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.013 | 0.28 | 0 | TN | ug N/L | 0.0703 | 1.08 | 0 |  | 1 | 1 |  | 0 | 0 | 0 | [6 empty] | 1 |  | 0 | 0 | 0 | [6 empty] | 24
METHOD CHECKS: | [41 empty] | 33
2026-02-27 | 12:38 | OM_2-27-2026_12-05-56PM.OMN | NO2 | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.137 | 1.72 | 37.5 | TN | ug N/L | 1.87 | 24.5 | 1500 | 102.449 | 1 | 1 |  | 37.5 | 37.5 | 37.5 | [6 empty] | 1 |  | 1500 | 1500 | 1500 | [6 empty] | 34
2026-02-27 | 12:39 | OM_2-27-2026_12-05-56PM.OMN | NO3 | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.138 | 1.7 | 37 | TN | ug N/L | 1.98 | 25.1 | 1540 |  | 1 | 1 |  | 37 | 37 | 37 | [6 empty] | 1 |  | 1540 | 1540 | 1540 | [6 empty] | 35
2026-02-27 | 12:41 | OM_2-27-2026_12-05-56PM.OMN | water | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.136 | 1.65 | 35.6 | TN | ug N/L | 0.00786 | 0.124 | -59.6 |  | 1 | 1 |  | 35.6 | 35.6 | 35.6 | [6 empty] | 1 |  | -59.6 | -59.6 | -59.6 | [6 empty] | 36
2026-02-27 | 12:42 | OM_2-27-2026_12-05-56PM.OMN | water blank | 19 | Unknown | 1 | 1 | TP | ug P/L | 0.0128 | 0.314 | -0.0151 | TN | ug N/L | 0.0703 | 1.1 | -0.261 |  | 1 | 1 |  | -0.0151 | -0.0151 | -0.0151 | [6 empty] | 1 |  | -0.261 | -0.261 | -0.261 | [6 empty] | 37
2026-02-27 | 12:43 | OM_2-27-2026_12-05-56PM.OMN | water blank | 20 | Unknown | 1 | 1 | TP | ug P/L | 0.0111 | 0.291 | -0.624 | TN | ug N/L | 0.0676 | 1.05 | -3.38 |  | 1 | 1 |  | -0.624 | -0.624 | -0.624 | [6 empty] | 1 |  | -3.38 | -3.38 | -3.38 | [6 empty] | 38
2026-02-27 | 12:45 | OM_2-27-2026_12-05-56PM.OMN | water blank | 21 | Unknown | 1 | 1 | TP | ug P/L | 0.0122 | 0.313 | -0.0603 | TN | ug N/L | 0.0688 | 1.06 | -2.77 |  | 1 | 1 |  | -0.0603 | -0.0603 | -0.0603 | [6 empty] | 1 |  | -2.77 | -2.77 | -2.77 | [6 empty] | 39
2026-02-27 | 12:46 | OM_2-27-2026_12-05-56PM.OMN | water blank | 22 | Unknown | 1 | 1 | TP | ug P/L | 0.0126 | 0.287 | -0.706 | TN | ug N/L | 0.069 | 1.06 | -2.88 |  | 1 | 1 |  | -0.706 | -0.706 | -0.706 | [6 empty] | 1 |  | -2.88 | -2.88 | -2.88 | [6 empty] | 40
2026-02-27 | 12:47 | OM_2-27-2026_12-05-56PM.OMN | water blank | 23 | Unknown | 1 | 1 | TP | ug P/L | 0.0124 | 0.312 | -0.0704 | TN | ug N/L | 0.0692 | 1.07 | -2.32 |  | 1 | 1 |  | -0.0704 | -0.0704 | -0.0704 | [6 empty] | 1 |  | -2.32 | -2.32 | -2.32 | [6 empty] | 41
2026-02-27 | 12:48 | OM_2-27-2026_12-05-56PM.OMN | digest check | 24 | Unknown | 1 | 1 | TP | ug P/L | 0.0782 | 1.64 | 35.2 | TN | ug N/L | 0.549 | 7.13 | 370 |  | 1 | 1 |  | 35.2 | 35.2 | 35.2 | [6 empty] | 1 |  | 370 | 370 | 370 | [6 empty] | 42
2026-02-27 | 12:50 | OM_2-27-2026_12-05-56PM.OMN | digest check | 25 | Unknown | 1 | 1 | TP | ug P/L | 0.0765 | 1.63 | 35 | TN | ug N/L | 0.57 | 7.4 | 387 |  | 1 | 1 |  | 35 | 35 | 35 | [6 empty] | 1 |  | 387 | 387 | 387 | [6 empty] | 43
2026-02-27 | 12:51 | OM_2-27-2026_12-05-56PM.OMN | digest check | 26 | Unknown | 1 | 1 | TP | ug P/L | 0.0756 | 1.63 | 35 | TN | ug N/L | 0.56 | 7.21 | 376 |  | 1 | 1 |  | 35 | 35 | 35 | [6 empty] | 1 |  | 376 | 376 | 376 | [6 empty] | 44
2026-02-27 | 12:52 | OM_2-27-2026_12-05-56PM.OMN | digest check | 27 | Unknown | 1 | 1 | TP | ug P/L | 0.086 | 1.7 | 36.9 | TN | ug N/L | 0.541 | 6.98 | 361 |  | 1 | 1 |  | 36.9 | 36.9 | 36.9 | [6 empty] | 1 |  | 361 | 361 | 361 | [6 empty] | 45
2026-02-27 | 12:53 | OM_2-27-2026_12-05-56PM.OMN | digest check | 28 | Unknown | 1 | 1 | TP | ug P/L | 0.0793 | 1.69 | 36.8 | TN | ug N/L | 0.562 | 7.28 | 380 |  | 1 | 1 |  | 36.8 | 36.8 | 36.8 | [6 empty] | 1 |  | 380 | 380 | 380 | [6 empty] | 46
2026-02-27 | 12:55 | OM_2-27-2026_12-05-56PM.OMN | spk blank | 29 | Unknown | 1 | 1 | TP | ug P/L | 0.051 | 1.09 | 20.2 | TN | ug N/L | 0.377 | 4.96 | 236 |  | 1 | 1 |  | 20.2 | 20.2 | 20.2 | [6 empty] | 1 |  | 236 | 236 | 236 | [6 empty] | 47
2026-02-27 | 12:56 | OM_2-27-2026_12-05-56PM.OMN | spk blank | 30 | Unknown | 1 | 1 | TP | ug P/L | 0.0516 | 1.12 | 21.1 | TN | ug N/L | 0.445 | 5.79 | 287 |  | 1 | 1 |  | 21.1 | 21.1 | 21.1 | [6 empty] | 1 |  | 287 | 287 | 287 | [6 empty] | 48
2026-02-27 | 12:57 | OM_2-27-2026_12-05-56PM.OMN | spk blank | 31 | Unknown | 1 | 1 | TP | ug P/L | 0.0519 | 1.15 | 21.8 | TN | ug N/L | 0.379 | 4.98 | 237 |  | 1 | 1 |  | 21.8 | 21.8 | 21.8 | [6 empty] | 1 |  | 237 | 237 | 237 | [6 empty] | 49
2026-02-27 | 12:59 | OM_2-27-2026_12-05-56PM.OMN | spk blank | 32 | Unknown | 1 | 1 | TP | ug P/L | 0.0488 | 1.06 | 19.6 | TN | ug N/L | 0.372 | 4.89 | 232 |  | 1 | 1 |  | 19.6 | 19.6 | 19.6 | [6 empty] | 1 |  | 232 | 232 | 232 | [6 empty] | 50
2026-02-27 | 13:00 | OM_2-27-2026_12-05-56PM.OMN | spk blank | 33 | Unknown | 1 | 1 | TP | ug P/L | 0.0488 | 1.07 | 19.7 | TN | ug N/L | 0.369 | 4.81 | 227 |  | 1 | 1 |  | 19.7 | 19.7 | 19.7 | [6 empty] | 1 |  | 227 | 227 | 227 | [6 empty] | 51
SAMPLES:
2026-02-27 | 13:01 | OM_2-27-2026_12-05-56PM.OMN | S1 | 34 | Unknown | 1 | 1 | TP | ug P/L | 0.523 | 9.65 | 307 | TN | ug N/L | 1.46 | 18.5 | 1100 |  | 1 | 3 | out of range | 307 | 307 | 307 | [6 empty] | 1 |  | 1100 | 1100 | 1100 | [6 empty] | 54
2026-02-27 | 13:02 | OM_2-27-2026_12-05-56PM.OMN | S2 | 35 | Unknown | 1 | 1 | TP | ug P/L | 0.146 | 2.84 | 69.7 | TN | ug N/L | 0.885 | 11.3 | 634 |  | 1 | 1 |  | 69.7 | 69.7 | 69.7 | [6 empty] | 1 |  | 634 | 634 | 634 | [6 empty] | 55
2026-02-27 | 13:04 | OM_2-27-2026_12-05-56PM.OMN | S3 | 36 | Unknown | 1 | 1 | TP | ug P/L | 0.854 | 15.5 | 568 | TN | ug N/L | 2.37 | 30.1 | 1880 |  | 1 | 3 | out of range | 568 | 568 | 568 | [6 empty] | 3 | out of range | 1880 | 1880 | 1880 | [6 empty] | 56
2026-02-27 | 13:05 | OM_2-27-2026_12-05-56PM.OMN | S4 | 37 | Unknown | 1 | 1 | TP | ug P/L | 0.0558 | 1.2 | 23.2 | TN | ug N/L | 1.13 | 14.4 | 834 |  | 1 | 1 |  | 23.2 | 23.2 | 23.2 | [6 empty] | 1 |  | 834 | 834 | 834 | [6 empty] | 57
2026-02-27 | 13:06 | OM_2-27-2026_12-05-56PM.OMN | S5 | 38 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.119 | 2.35 | 55.3 | TN | ug N/L | 1.16 | 14.8 | 861 |  | 1 | 1 |  | 55.3 | 55.3 | 55.3 | [6 empty] | 1 |  | 861 | 861 | 861 | [6 empty] | 58
2026-02-27 | 13:07 | OM_2-27-2026_12-05-56PM.OMN | DUP S5 | 44 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.112 | 2.44 | 57.9 | TN | ug N/L | 1.12 | 14.3 | 825 |  | 1 | 1 |  | 57.9 | 57.9 | 57.9 |  |  | 4.5936 | PASS |  |  | 1 |  | 825 | 825 | 825 |  |  | 4.2705 | PASS |  |  | 59
2026-02-27 | 13:09 | OM_2-27-2026_12-05-56PM.OMN | S6 | 39 | Unknown | 1 | 1 | TP | ug P/L | 0.226 | 4.31 | 115 | TN | ug N/L | 1.34 | 17.1 | 1000 |  | 1 | 1 |  | 115 | 115 | 115 | [6 empty] | 1 |  | 1000 | 1000 | 1000 | [6 empty] | 60
2026-02-27 | 13:10 | OM_2-27-2026_12-05-56PM.OMN | S7 | 40 | Unknown | 1 | 1 | TP | ug P/L | 0.14 | 2.74 | 66.8 | TN | ug N/L | 1 | 12.9 | 733 |  | 1 | 1 |  | 66.8 | 66.8 | 66.8 | [6 empty] | 1 |  | 733 | 733 | 733 | [6 empty] | 61
2026-02-27 | 13:11 | OM_2-27-2026_12-05-56PM.OMN | S8 | 41 | Unknown | 1 | 1 | TP | ug P/L | 0.0824 | 1.79 | 39.5 | TN | ug N/L | 0.662 | 8.44 | 453 |  | 1 | 1 |  | 39.5 | 39.5 | 39.5 | [6 empty] | 1 |  | 453 | 453 | 453 | [6 empty] | 62
2026-02-27 | 13:13 | OM_2-27-2026_12-05-56PM.OMN | S9 | 42 | Unknown | 1 | 1 | TP | ug P/L | 0.0595 | 1.31 | 26.2 | TN | ug N/L | 0.547 | 7.04 | 365 |  | 1 | 1 |  | 26.2 | 26.2 | 26.2 | [6 empty] | 1 |  | 365 | 365 | 365 | [6 empty] | 63
2026-02-27 | 13:14 | OM_2-27-2026_12-05-56PM.OMN | S10 | 43 | Spike Set | 1 | 1 | TP | ug P/L | 0.955 | 16.3 | 606 | TN | ug N/L | 1.14 | 14.2 | 816 |  | 1 | 3 | out of range | 606 | 606 | 606 | [6 empty] | 1 |  | 816 | 816 | 816 | [6 empty] | 64
2026-02-27 | 13:15 | OM_2-27-2026_12-05-56PM.OMN | SPK S10 | 45 | Spike Set | 1 | 1 | TP | ug P/L | 0.923 | 16.7 | 627 | TN | ug N/L | 1.33 | 16.7 | 983 |  | 1 | 3 | out of range | 627 | 627 | 627 |  |  |  |  | 146.7 | FAIL | 1 |  | 983 | 983 | 983 |  |  |  |  | 76.63 | FAIL | 65
2026-02-27 | 13:16 | OM_2-27-2026_12-05-56PM.OMN | water | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.139 | 1.51 | 31.7 | TN | ug N/L | 0.00601 | 0.0787 | -62.4 |  | 1 | 1 |  | 31.7 | 31.7 | 31.7 | [6 empty] | 1 |  | -62.4 | -62.4 | -62.4 | [6 empty] | 66
2026-02-27 | 13:18 | OM_2-27-2026_12-05-56PM.OMN | CHECK25 | 15 | Check Standard | 1 | 1 | TP | ug P/L | 0.0521 | 1.17 | 22.6 | TN | ug N/L | 0.366 | 4.8 | 226 |  | 1 | 1 |  | 22.6 | 22.6 | 22.6 | -9.6 | PASS |  |  |  |  | 1 |  | 226 | 226 | 226 | -9.6 | PASS |  |  |  |  | 67
2026-02-27 | 13:19 | OM_2-27-2026_12-05-56PM.OMN | S11 | 46 | Unknown | 1 | 1 | TP | ug P/L | 0.162 | 3.08 | 76.7 | TN | ug N/L | 0.354 | 4.6 | 214 |  | 1 | 1 |  | 76.7 | 76.7 | 76.7 | [6 empty] | 1 |  | 214 | 214 | 214 | [6 empty] | 68
2026-02-27 | 13:20 | OM_2-27-2026_12-05-56PM.OMN | S12 | 47 | Unknown | 1 | 1 | TP | ug P/L | 1.04 | 18.7 | 736 | TN | ug N/L | 2.24 | 28.2 | 1750 |  | 1 | 3 | out of range | 736 | 736 | 736 | [6 empty] | 3 | out of range | 1750 | 1750 | 1750 | [6 empty] | 69
2026-02-27 | 13:22 | OM_2-27-2026_12-05-56PM.OMN | S13 | 48 | Unknown | 1 | 1 | TP | ug P/L | 0.11 | 2.21 | 51.4 | TN | ug N/L | 1.35 | 17.1 | 1010 |  | 1 | 1 |  | 51.4 | 51.4 | 51.4 | [6 empty] | 1 |  | 1010 | 1010 | 1010 | [6 empty] | 70
2026-02-27 | 13:23 | OM_2-27-2026_12-05-56PM.OMN | S14 | 49 | Unknown | 1 | 1 | TP | ug P/L | 0.0914 | 1.73 | 37.7 | TN | ug N/L | 0.33 | 4.33 | 197 |  | 1 | 1 |  | 37.7 | 37.7 | 37.7 | [6 empty] | 1 |  | 197 | 197 | 197 | [6 empty] | 71
2026-02-27 | 13:24 | OM_2-27-2026_12-05-56PM.OMN | S15 | 50 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.163 | 3.18 | 79.8 | TN | ug N/L | 1.23 | 15.6 | 910 |  | 1 | 1 |  | 79.8 | 79.8 | 79.8 | [6 empty] | 1 |  | 910 | 910 | 910 | [6 empty] | 72
2026-02-27 | 13:25 | OM_2-27-2026_12-05-56PM.OMN | DUP S15 | 56 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.171 | 3.36 | 85.3 | TN | ug N/L | 1.24 | 15.9 | 929 |  | 1 | 1 |  | 85.3 | 85.3 | 85.3 |  |  | 6.6626 | PASS |  |  | 1 |  | 929 | 929 | 929 |  |  | 2.0663 | PASS |  |  | 73
2026-02-27 | 13:27 | OM_2-27-2026_12-05-56PM.OMN | S16 | 51 | Unknown | 1 | 1 | TP | ug P/L | 0.0577 | 1.27 | 25.2 | TN | ug N/L | 0.757 | 9.69 | 532 |  | 1 | 1 |  | 25.2 | 25.2 | 25.2 | [6 empty] | 1 |  | 532 | 532 | 532 | [6 empty] | 74
2026-02-27 | 13:28 | OM_2-27-2026_12-05-56PM.OMN | S17 | 52 | Unknown | 1 | 1 | TP | ug P/L | 0.116 | 2.28 | 53.4 | TN | ug N/L | 1.28 | 16.3 | 953 |  | 1 | 1 |  | 53.4 | 53.4 | 53.4 | [6 empty] | 1 |  | 953 | 953 | 953 | [6 empty] | 75
2026-02-27 | 13:29 | OM_2-27-2026_12-05-56PM.OMN | S18 | 53 | Unknown | 1 | 1 | TP | ug P/L | 0.0692 | 1.43 | 29.5 | TN | ug N/L | 0.659 | 8.46 | 454 |  | 1 | 1 |  | 29.5 | 29.5 | 29.5 | [6 empty] | 1 |  | 454 | 454 | 454 | [6 empty] | 76
2026-02-27 | 13:30 | OM_2-27-2026_12-05-56PM.OMN | S19 | 54 | Unknown | 1 | 1 | TP | ug P/L | 0.226 | 4.29 | 114 | TN | ug N/L | 1.37 | 17.4 | 1020 |  | 1 | 1 |  | 114 | 114 | 114 | [6 empty] | 1 |  | 1020 | 1020 | 1020 | [6 empty] | 77
2026-02-27 | 13:32 | OM_2-27-2026_12-05-56PM.OMN | S20 | 55 | Spike Set | 1 | 1 | TP | ug P/L | 0.102 | 2.08 | 47.6 | TN | ug N/L | 0.67 | 8.75 | 472 |  | 1 | 1 |  | 47.6 | 47.6 | 47.6 | [6 empty] | 1 |  | 472 | 472 | 472 | [6 empty] | 78
2026-02-27 | 13:33 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S20 | 57 | Spike Set | 1 | 1 | TP | ug P/L | 0.149 | 2.89 | 71.1 | TN | ug N/L | 0.947 | 12.1 | 685 |  | 1 | 1 |  | 71.1 | 71.1 | 71.1 |  |  |  |  | 101.11 | PASS | 1 |  | 685 | 685 | 685 |  |  |  |  | 92.05 | PASS | 79
2026-02-27 | 13:34 | OM_2-27-2026_12-05-56PM.OMN | bisulfate | S4 | Unknown | 1 | 1 | TP | ug P/L | 0.00835 | 0.0885 | -5.76 | TN | ug N/L | 0.0202 | 0.397 | -43.1 |  | 1 | 1 |  | -5.76 | -5.76 | -5.76 | [6 empty] | 1 |  | -43.1 | -43.1 | -43.1 | [6 empty] | 80
2026-02-27 | 13:36 | OM_2-27-2026_12-05-56PM.OMN | CHECK25 | 16 | Check Standard | 1 | 1 | TP | ug P/L | 0.0593 | 1.28 | 25.4 | TN | ug N/L | 0.435 | 5.62 | 277 |  | 1 | 1 |  | 25.4 | 25.4 | 25.4 | 1.6 | PASS |  |  |  |  | 1 |  | 277 | 277 | 277 | 10.8 | PASS |  |  |  |  | 81
2026-02-27 | 13:37 | OM_2-27-2026_12-05-56PM.OMN | NO2 | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.138 | 1.73 | 37.9 | TN | ug N/L | 1.94 | 25.1 | 1540 | 102.7888 | 1 | 1 |  | 37.9 | 37.9 | 37.9 | [6 empty] | 1 |  | 1540 | 1540 | 1540 | [6 empty] | 82
2026-02-27 | 13:38 | OM_2-27-2026_12-05-56PM.OMN | NO3 | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.135 | 1.68 | 36.5 | TN | ug N/L | 2.03 | 25.8 | 1580 |  | 1 | 1 |  | 36.5 | 36.5 | 36.5 | [6 empty] | 1 |  | 1580 | 1580 | 1580 | [6 empty] | 83
2026-02-27 | 13:39 | OM_2-27-2026_12-05-56PM.OMN | S21 | 58 | Unknown | 1 | 1 | TP | ug P/L | 0.374 | 6.95 | 204 | TN | ug N/L | 0.582 | 7.64 | 402 |  | 1 | 3 | out of range | 204 | 204 | 204 | [6 empty] | 1 |  | 402 | 402 | 402 | [6 empty] | 84
2026-02-27 | 13:41 | OM_2-27-2026_12-05-56PM.OMN | S22 | 59 | Unknown | 1 | 1 | TP | ug P/L | 0.495 | 9.05 | 283 | TN | ug N/L | 1.51 | 15.5 | 902 |  | 1 | 3 | out of range | 283 | 283 | 283 | [6 empty] | 1 |  | 902 | 902 | 902 | [6 empty] | 85
2026-02-27 | 13:42 | OM_2-27-2026_12-05-56PM.OMN | S23 | 60 | Unknown | 1 | 1 | TP | ug P/L | 0.198 | 3.77 | 97.8 | TN | ug N/L | 0.946 | 12.1 | 686 |  | 1 | 1 |  | 97.8 | 97.8 | 97.8 | [6 empty] | 1 |  | 686 | 686 | 686 | [6 empty] | 86
2026-02-27 | 13:43 | OM_2-27-2026_12-05-56PM.OMN | S24 | 61 | Unknown | 1 | 1 | TP | ug P/L | 0.296 | 5.56 | 156 | TN | ug N/L | 3.25 | 41.4 | 2680 |  | 1 | 1 |  | 156 | 156 | 156 | [6 empty] | 3 | out of range | 2680 | 2680 | 2680 | [6 empty] | 87
2026-02-27 | 13:45 | OM_2-27-2026_12-05-56PM.OMN | S25 | 62 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.178 | 3.38 | 86 | TN | ug N/L | 1.24 | 15.7 | 913 |  | 1 | 1 |  | 86 | 86 | 86 | [6 empty] | 1 |  | 913 | 913 | 913 | [6 empty] | 88
2026-02-27 | 13:46 | OM_2-27-2026_12-05-56PM.OMN | DUP S25 | 68 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.174 | 3.36 | 85.3 | TN | ug N/L | 1.23 | 15.5 | 905 |  | 1 | 1 |  | 85.3 | 85.3 | 85.3 |  |  | 0.81728 | PASS |  |  | 1 |  | 905 | 905 | 905 |  |  | 0.88009 | PASS |  |  | 89
2026-02-27 | 13:47 | OM_2-27-2026_12-05-56PM.OMN | S26 | 63 | Unknown | 1 | 1 | TP | ug P/L | 0.348 | 6.47 | 187 | TN | ug N/L | 2.09 | 26.5 | 1630 |  | 1 | 3 | out of range | 187 | 187 | 187 | [6 empty] | 1 |  | 1630 | 1630 | 1630 | [6 empty] | 90
2026-02-27 | 13:48 | OM_2-27-2026_12-05-56PM.OMN | S27 | 64 | Unknown | 1 | 1 | TP | ug P/L | 0.106 | 2.14 | 49.5 | TN | ug N/L | 0.587 | 7.55 | 397 |  | 1 | 1 |  | 49.5 | 49.5 | 49.5 | [6 empty] | 1 |  | 397 | 397 | 397 | [6 empty] | 91
2026-02-27 | 13:50 | OM_2-27-2026_12-05-56PM.OMN | S28 | 65 | Unknown | 1 | 1 | TP | ug P/L | 0.0996 | 2.06 | 47.1 | TN | ug N/L | 1.16 | 14.8 | 858 |  | 1 | 1 |  | 47.1 | 47.1 | 47.1 | [6 empty] | 1 |  | 858 | 858 | 858 | [6 empty] | 92
2026-02-27 | 13:51 | OM_2-27-2026_12-05-56PM.OMN | S29 | 66 | Unknown | 1 | 1 | TP | ug P/L | 0.113 | 2.24 | 52.1 | TN | ug N/L | 1.11 | 14.1 | 815 |  | 1 | 1 |  | 52.1 | 52.1 | 52.1 | [6 empty] | 1 |  | 815 | 815 | 815 | [6 empty] | 93
2026-02-27 | 13:52 | OM_2-27-2026_12-05-56PM.OMN | S30 | 67 | Spike Set | 1 | 1 | TP | ug P/L | 0.0637 | 1.35 | 27.3 | TN | ug N/L | 0.976 | 12.5 | 709 |  | 1 | 1 |  | 27.3 | 27.3 | 27.3 | [6 empty] | 1 |  | 709 | 709 | 709 | [6 empty] | 94
2026-02-27 | 13:53 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S30 | 69 | Spike Set | 1 | 1 | TP | ug P/L | 0.13 | 2.34 | 55.1 | TN | ug N/L | 1.48 | 18.7 | 1110 |  | 1 | 1 |  | 55.1 | 55.1 | 55.1 |  |  |  |  | 116.71 | PASS | 1 |  | 1110 | 1110 | 1110 |  |  |  |  | 171.5 | FAIL | 95
2026-02-27 | 13:55 | OM_2-27-2026_12-05-56PM.OMN | water | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.157 | 1.89 | 42.2 | TN | ug N/L | 0.00555 | -0.115 | -74 |  | 1 | 3 | messy peak integration | 42.2 | 42.2 | 42.2 | [6 empty] | 1 |  | -74 | -74 | -74 | [6 empty] | 96
2026-02-27 | 13:56 | OM_2-27-2026_12-05-56PM.OMN | CHECK25 | 17 | Check Standard | 1 | 1 | TP | ug P/L | 0.0509 | 1.16 | 22.1 | TN | ug N/L | 0.374 | 4.88 | 231 |  | 1 | 1 |  | 22.1 | 22.1 | 22.1 | -11.6 | PASS |  |  |  |  | 1 |  | 231 | 231 | 231 | -7.6 | PASS |  |  |  |  | 97
2026-02-27 | 13:57 | OM_2-27-2026_12-05-56PM.OMN | S31 | 70 | Unknown | 1 | 1 | TP | ug P/L | 0.0914 | 1.86 | 41.4 | TN | ug N/L | 0.809 | 10.4 | 575 |  | 1 | 1 |  | 41.4 | 41.4 | 41.4 | [6 empty] | 1 |  | 575 | 575 | 575 | [6 empty] | 98
2026-02-27 | 13:59 | OM_2-27-2026_12-05-56PM.OMN | S32 | 71 | Unknown | 1 | 1 | TP | ug P/L | 0.487 | 8.97 | 280 | TN | ug N/L | 0.756 | 9.56 | 523 |  | 1 | 3 | out of range | 280 | 280 | 280 | [6 empty] | 1 |  | 523 | 523 | 523 | [6 empty] | 99
2026-02-27 | 14:00 | OM_2-27-2026_12-05-56PM.OMN | S33 | 72 | Unknown | 1 | 1 | TP | ug P/L | 0.0946 | 1.95 | 43.9 | TN | ug N/L | 0.9 | 11.4 | 641 |  | 1 | 1 |  | 43.9 | 43.9 | 43.9 | [6 empty] | 1 |  | 641 | 641 | 641 | [6 empty] | 100
2026-02-27 | 14:01 | OM_2-27-2026_12-05-56PM.OMN | S34 | 73 | Unknown | 1 | 1 | TP | ug P/L | 0.102 | 2.04 | 46.5 | TN | ug N/L | 1.85 | 23.5 | 1430 |  | 1 | 1 |  | 46.5 | 46.5 | 46.5 | [6 empty] | 1 |  | 1430 | 1430 | 1430 | [6 empty] | 101
2026-02-27 | 14:02 | OM_2-27-2026_12-05-56PM.OMN | S35 | 74 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.242 | 4.61 | 124 | TN | ug N/L | 1.12 | 14.2 | 821 |  | 1 | 1 |  | 124 | 124 | 124 | [6 empty] | 1 |  | 821 | 821 | 821 | [6 empty] | 102
2026-02-27 | 14:04 | OM_2-27-2026_12-05-56PM.OMN | DUP S35 | 80 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.251 | 4.74 | 129 | TN | ug N/L | 1.16 | 14.7 | 853 |  | 1 | 1 |  | 129 | 129 | 129 |  |  | 3.9526 | PASS |  |  | 1 |  | 853 | 853 | 853 |  |  | 3.8232 | PASS |  |  | 103
2026-02-27 | 14:05 | OM_2-27-2026_12-05-56PM.OMN | S36 | 75 | Unknown | 1 | 1 | TP | ug P/L | 0.0883 | 1.84 | 40.8 | TN | ug N/L | 0.923 | 11.8 | 665 |  | 1 | 1 |  | 40.8 | 40.8 | 40.8 | [6 empty] | 1 |  | 665 | 665 | 665 | [6 empty] | 104
2026-02-27 | 14:06 | OM_2-27-2026_12-05-56PM.OMN | S37 | 76 | Unknown | 1 | 1 | TP | ug P/L | 0.614 | 11.3 | 374 | TN | ug N/L | 1.51 | 19.2 | 1140 |  | 1 | 3 | out of range | 374 | 374 | 374 | [6 empty] | 1 |  | 1140 | 1140 | 1140 | [6 empty] | 105
2026-02-27 | 14:08 | OM_2-27-2026_12-05-56PM.OMN | S38 | 77 | Unknown | 1 | 1 | TP | ug P/L | 0.104 | 2.08 | 47.7 | TN | ug N/L | 0.799 | 10.2 | 566 |  | 1 | 1 |  | 47.7 | 47.7 | 47.7 | [6 empty] | 1 |  | 566 | 566 | 566 | [6 empty] | 106
2026-02-27 | 14:09 | OM_2-27-2026_12-05-56PM.OMN | S39 | 78 | Unknown | 1 | 1 | TP | ug P/L | 0.859 | 15.8 | 583 | TN | ug N/L | 3.17 | 40 | 2580 |  | 1 | 3 | out of range | 583 | 583 | 583 | [6 empty] | 3 | out of range | 2580 | 2580 | 2580 | [6 empty] | 107
2026-02-27 | 14:10 | OM_2-27-2026_12-05-56PM.OMN | S40 | 79 | Spike Set | 1 | 1 | TP | ug P/L | 0.0597 | 1.3 | 26.1 | TN | ug N/L | 0.723 | 9.28 | 505 |  | 1 | 1 |  | 26.1 | 26.1 | 26.1 | [6 empty] | 1 |  | 505 | 505 | 505 | [6 empty] | 108
2026-02-27 | 14:11 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S40 | 81 | Spike Set | 1 | 1 | TP | ug P/L | 0.106 | 2.14 | 49.3 | TN | ug N/L | 0.973 | 12.4 | 704 |  | 1 | 1 |  | 49.3 | 49.3 | 49.3 |  |  |  |  | 97.73 | PASS | 1 |  | 704 | 704 | 704 |  |  |  |  | 86.64 | PASS | 109
2026-02-27 | 14:13 | OM_2-27-2026_12-05-56PM.OMN | BISULFATE | S4 | Unknown | 1 | 1 | TP | ug P/L | 0.00831 | 0.083 | -5.9 | TN | ug N/L | 0.0192 | 0.383 | -43.9 |  | 1 | 1 |  | -5.9 | -5.9 | -5.9 | [6 empty] | 1 |  | -43.9 | -43.9 | -43.9 | [6 empty] | 110
2026-02-27 | 14:14 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 | 18 | Check Standard | 1 | 1 | TP | ug P/L | 0.0511 | 1.13 | 21.5 | TN | ug N/L | 0.376 | 4.97 | 236 |  | 1 | 1 |  | 21.5 | 21.5 | 21.5 | -14 | PASS |  |  |  |  | 1 |  | 236 | 236 | 236 | -5.6 | PASS |  |  |  |  | 111
2026-02-27 | 14:15 | OM_2-27-2026_12-05-56PM.OMN | NO2 | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.135 | 1.64 | 35.4 | TN | ug N/L | 1.83 | 24 | 1460 | 104.1667 | 1 | 1 |  | 35.4 | 35.4 | 35.4 | [6 empty] | 1 |  | 1460 | 1460 | 1460 | [6 empty] | 112
2026-02-27 | 14:17 | OM_2-27-2026_12-05-56PM.OMN | NO3 | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.139 | 1.72 | 37.5 | TN | ug N/L | 1.97 | 25 | 1530 |  | 1 | 1 |  | 37.5 | 37.5 | 37.5 | [6 empty] | 1 |  | 1530 | 1530 | 1530 | [6 empty] | 113
2026-02-27 | 14:18 | OM_2-27-2026_12-05-56PM.OMN | S41 | 82 | Unknown | 1 | 1 | TP | ug P/L | 0.106 | 2.16 | 49.8 | TN | ug N/L | 0.823 | 10.6 | 591 |  | 1 | 1 |  | 49.8 | 49.8 | 49.8 | [6 empty] | 1 |  | 591 | 591 | 591 | [6 empty] | 114
2026-02-27 | 14:19 | OM_2-27-2026_12-05-56PM.OMN | S42 | 83 | Unknown | 1 | 1 | TP | ug P/L | 0.216 | 3.82 | 99.3 | TN | ug N/L | 0.346 | 3.98 | 176 |  | 1 | 1 |  | 99.3 | 99.3 | 99.3 | [6 empty] | 1 |  | 176 | 176 | 176 | [6 empty] | 115
2026-02-27 | 14:20 | OM_2-27-2026_12-05-56PM.OMN | S43 | 84 | Unknown | 1 | 1 | TP | ug P/L | 0.378 | 6.81 | 199 | TN | ug N/L | 2.86 | 36.3 | 2310 |  | 1 | 3 | out of range | 199 | 199 | 199 | [6 empty] | 3 | out of range | 2310 | 2310 | 2310 | [6 empty] | 116
2026-02-27 | 14:22 | OM_2-27-2026_12-05-56PM.OMN | S44 | 85 | Unknown | 1 | 1 | TP | ug P/L | 0.0606 | 1.31 | 26.3 | TN | ug N/L | 0.541 | 7 | 363 |  | 1 | 1 |  | 26.3 | 26.3 | 26.3 | [6 empty] | 1 |  | 363 | 363 | 363 | [6 empty] | 117
2026-02-27 | 14:23 | OM_2-27-2026_12-05-56PM.OMN | S45 | 86 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.103 | 2.05 | 46.8 | TN | ug N/L | 0.956 | 12.2 | 692 |  | 1 | 1 |  | 46.8 | 46.8 | 46.8 | [6 empty] | 1 |  | 692 | 692 | 692 | [6 empty] | 118
2026-02-27 | 14:24 | OM_2-27-2026_12-05-56PM.OMN | DUP S45 | 92 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.101 | 2.01 | 45.8 | TN | ug N/L | 1.06 | 13.6 | 778 |  | 1 | 1 |  | 45.8 | 45.8 | 45.8 |  |  | 2.1598 | PASS |  |  | 1 |  | 778 | 778 | 778 |  |  | 11.7007 | PASS |  |  | 119
2026-02-27 | 14:25 | OM_2-27-2026_12-05-56PM.OMN | S46 | 87 | Unknown | 1 | 1 | TP | ug P/L | 0.287 | 5.33 | 148 | TN | ug N/L | 1.48 | 18.8 | 1120 |  | 1 | 1 |  | 148 | 148 | 148 | [6 empty] | 1 |  | 1120 | 1120 | 1120 | [6 empty] | 120
2026-02-27 | 14:27 | OM_2-27-2026_12-05-56PM.OMN | S47 | 88 | Unknown | 1 | 1 | TP | ug P/L | 0.119 | 2.39 | 56.6 | TN | ug N/L | 0.381 | 4.97 | 237 |  | 1 | 1 |  | 56.6 | 56.6 | 56.6 | [6 empty] | 1 |  | 237 | 237 | 237 | [6 empty] | 121
2026-02-27 | 14:28 | OM_2-27-2026_12-05-56PM.OMN | S48 | 89 | Unknown | 1 | 1 | TP | ug P/L | 0.111 | 2.22 | 51.6 | TN | ug N/L | 0.814 | 10.4 | 577 |  | 1 | 1 |  | 51.6 | 51.6 | 51.6 | [6 empty] | 1 |  | 577 | 577 | 577 | [6 empty] | 122
2026-02-27 | 14:29 | OM_2-27-2026_12-05-56PM.OMN | S49 | 90 | Unknown | 1 | 1 | TP | ug P/L | 0.166 | 2.71 | 65.8 | TN | ug N/L | 1.58 | 20.1 | 1210 |  | 1 | 1 |  | 65.8 | 65.8 | 65.8 | [6 empty] | 1 |  | 1210 | 1210 | 1210 | [6 empty] | 123
2026-02-27 | 14:31 | OM_2-27-2026_12-05-56PM.OMN | S50 | 91 | Spike Set | 1 | 1 | TP | ug P/L | 0.13 | 2.57 | 61.7 | TN | ug N/L | 1.05 | 13.4 | 769 |  | 1 | 1 |  | 61.7 | 61.7 | 61.7 | [6 empty] | 1 |  | 769 | 769 | 769 | [6 empty] | 124
2026-02-27 | 14:32 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S50 | 93 | Spike Set | 1 | 1 | TP | ug P/L | 0.178 | 3.35 | 85 | TN | ug N/L | 1.32 | 16.8 | 985 |  | 1 | 1 |  | 85 | 85 | 85 |  |  |  |  | 101.7 | PASS | 1 |  | 985 | 985 | 985 |  |  |  |  | 96.25 | PASS | 125
2026-02-27 | 14:33 | OM_2-27-2026_12-05-56PM.OMN | BLANK | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.135 | 1.63 | 35 | TN | ug N/L | 0.00585 | 0.114 | -60.2 |  | 1 | 1 |  | 35 | 35 | 35 | [6 empty] | 1 |  | -60.2 | -60.2 | -60.2 | [6 empty] | 126
2026-02-27 | 14:34 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 | 15 | Check Standard | 1 | 1 | TP | ug P/L | 0.0518 | 1.12 | 21.2 | TN | ug N/L | 0.398 | 4.92 | 233 |  | 1 | 1 |  | 21.2 | 21.2 | 21.2 | -15.2 | PASS |  |  |  |  | 1 |  | 233 | 233 | 233 | -6.8 | PASS |  |  |  |  | 127
2026-02-27 | 14:36 | OM_2-27-2026_12-05-56PM.OMN | S51 | 94 | Unknown | 1 | 1 | TP | ug P/L | 0.0488 | 1.14 | 21.6 | TN | ug N/L | 0.564 | 7.17 | 373 |  | 1 | 1 |  | 21.6 | 21.6 | 21.6 | [6 empty] | 1 |  | 373 | 373 | 373 | [6 empty] | 128
2026-02-27 | 14:37 | OM_2-27-2026_12-05-56PM.OMN | S52 | 95 | Unknown | 1 | 1 | TP | ug P/L | 0.157 | 3.03 | 75.3 | TN | ug N/L | 0.957 | 12.2 | 693 |  | 1 | 1 |  | 75.3 | 75.3 | 75.3 | [6 empty] | 1 |  | 693 | 693 | 693 | [6 empty] | 129
2026-02-27 | 14:38 | OM_2-27-2026_12-05-56PM.OMN | S53 | 96 | Unknown | 1 | 1 | TP | ug P/L | 0.282 | 5.28 | 146 | TN | ug N/L | 1.58 | 20 | 1200 |  | 1 | 1 |  | 146 | 146 | 146 | [6 empty] | 1 |  | 1200 | 1200 | 1200 | [6 empty] | 130
2026-02-27 | 14:40 | OM_2-27-2026_12-05-56PM.OMN | S54 | 97 | Unknown | 1 | 1 | TP | ug P/L | 0.23 | 4.4 | 117 | TN | ug N/L | 0.35 | 4.53 | 210 |  | 1 | 1 |  | 117 | 117 | 117 | [6 empty] | 1 |  | 210 | 210 | 210 | [6 empty] | 131
2026-02-27 | 14:41 | OM_2-27-2026_12-05-56PM.OMN | S55 | 98 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.0587 | 1.26 | 24.9 | TN | ug N/L | 0.629 | 8.15 | 434 |  | 1 | 1 |  | 24.9 | 24.9 | 24.9 | [6 empty] | 1 |  | 434 | 434 | 434 | [6 empty] | 132
2026-02-27 | 14:42 | OM_2-27-2026_12-05-56PM.OMN | DUP S55 | 104 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.0553 | 1.2 | 23.2 | TN | ug N/L | 0.678 | 8.31 | 444 |  | 1 | 1 |  | 23.2 | 23.2 | 23.2 |  |  | 7.0686 | PASS |  |  | 1 |  | 444 | 444 | 444 |  |  | 2.2779 | PASS |  |  | 133
2026-02-27 | 14:43 | OM_2-27-2026_12-05-56PM.OMN | S56 | 99 | Unknown | 1 | 1 | TP | ug P/L | 0.0578 | 1.27 | 25.2 | TN | ug N/L | 0.963 | 12.2 | 692 |  | 1 | 1 |  | 25.2 | 25.2 | 25.2 | [6 empty] | 1 |  | 692 | 692 | 692 | [6 empty] | 134
2026-02-27 | 14:45 | OM_2-27-2026_12-05-56PM.OMN | S57 | 100 | Unknown | 1 | 1 | TP | ug P/L | 0.143 | 2.73 | 66.5 | TN | ug N/L | 1.16 | 14.7 | 851 |  | 1 | 1 |  | 66.5 | 66.5 | 66.5 | [6 empty] | 1 |  | 851 | 851 | 851 | [6 empty] | 135
2026-02-27 | 14:46 | OM_2-27-2026_12-05-56PM.OMN | S58 | 101 | Unknown | 1 | 1 | TP | ug P/L | 0.111 | 2.27 | 53 | TN | ug N/L | 0.682 | 8.77 | 474 |  | 1 | 1 |  | 53 | 53 | 53 | [6 empty] | 1 |  | 474 | 474 | 474 | [6 empty] | 136
2026-02-27 | 14:47 | OM_2-27-2026_12-05-56PM.OMN | S59 | 102 | Unknown | 1 | 1 | TP | ug P/L | 0.0892 | 1.82 | 40.4 | TN | ug N/L | 0.94 | 12 | 679 |  | 1 | 1 |  | 40.4 | 40.4 | 40.4 | [6 empty] | 1 |  | 679 | 679 | 679 | [6 empty] | 137
2026-02-27 | 14:48 | OM_2-27-2026_12-05-56PM.OMN | S60 | 103 | Spike Set | 1 | 1 | TP | ug P/L | 0.0639 | 1.37 | 28 | TN | ug N/L | 0.604 | 7.78 | 411 |  | 1 | 1 |  | 28 | 28 | 28 | [6 empty] | 1 |  | 411 | 411 | 411 | [6 empty] | 138
2026-02-27 | 14:50 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S60 | 105 | Spike Set | 1 | 1 | TP | ug P/L | 0.11 | 2.26 | 52.7 | TN | ug N/L | 0.892 | 11.4 | 638 |  | 1 | 1 |  | 52.7 | 52.7 | 52.7 |  |  |  |  | 104.07 | PASS | 1 |  | 638 | 638 | 638 |  |  |  |  | 97.18 | PASS | 139
2026-02-27 | 14:51 | OM_2-27-2026_12-05-56PM.OMN | BISULFATE | S4 | Unknown | 1 | 1 | TP | ug P/L | 0.0178 | 0.216 | -2.53 | TN | ug N/L | 0.0165 | 0.244 | -52.4 |  | 1 | 1 |  | -2.53 | -2.53 | -2.53 | [6 empty] | 1 |  | -52.4 | -52.4 | -52.4 | [6 empty] | 140
2026-02-27 | 14:52 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 | 16 | Check Standard | 1 | 1 | TP | ug P/L | 0.0503 | 1.14 | 21.6 | TN | ug N/L | 0.371 | 4.84 | 229 |  | 1 | 1 |  | 21.6 | 21.6 | 21.6 | -13.6 | PASS |  |  |  |  | 1 |  | 229 | 229 | 229 | -8.4 | PASS |  |  |  |  | 141
2026-02-27 | 14:54 | OM_2-27-2026_12-05-56PM.OMN | NO2 | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.134 | 1.63 | 35.1 | TN | ug N/L | 1.87 | 24.4 | 1490 | 100.8197 | 1 | 1 |  | 35.1 | 35.1 | 35.1 | [6 empty] | 1 |  | 1490 | 1490 | 1490 | [6 empty] | 142
2026-02-27 | 14:55 | OM_2-27-2026_12-05-56PM.OMN | NO3 | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.135 | 1.61 | 34.6 | TN | ug N/L | 1.94 | 24.6 | 1510 |  | 1 | 1 |  | 34.6 | 34.6 | 34.6 | [6 empty] | 1 |  | 1510 | 1510 | 1510 | [6 empty] | 143
2026-02-27 | 14:56 | OM_2-27-2026_12-05-56PM.OMN | S61 | 106 | Unknown | 1 | 1 | TP | ug P/L | 0.145 | 2.84 | 69.6 | TN | ug N/L | 1.03 | 13.1 | 748 |  | 1 | 1 |  | 69.6 | 69.6 | 69.6 | [6 empty] | 1 |  | 748 | 748 | 748 | [6 empty] | 144
2026-02-27 | 14:57 | OM_2-27-2026_12-05-56PM.OMN | S62 | 107 | Unknown | 1 | 1 | TP | ug P/L | 0.0886 | 1.8 | 39.7 | TN | ug N/L | 0.373 | 4.88 | 231 |  | 1 | 1 |  | 39.7 | 39.7 | 39.7 | [6 empty] | 1 |  | 231 | 231 | 231 | [6 empty] | 145
2026-02-27 | 14:59 | OM_2-27-2026_12-05-56PM.OMN | S63 | 108 | Unknown | 1 | 1 | TP | ug P/L | 0.101 | 1.91 | 42.8 | TN | ug N/L | 1.01 | 12.8 | 732 |  | 1 | 1 |  | 42.8 | 42.8 | 42.8 | [6 empty] | 1 |  | 732 | 732 | 732 | [6 empty] | 146
2026-02-27 | 15:00 | OM_2-27-2026_12-05-56PM.OMN | S64 | 109 | Unknown | 1 | 1 | TP | ug P/L | 0.0809 | 1.71 | 37.3 | TN | ug N/L | 0.967 | 12.3 | 698 |  | 1 | 1 |  | 37.3 | 37.3 | 37.3 | [6 empty] | 1 |  | 698 | 698 | 698 | [6 empty] | 147
2026-02-27 | 15:01 | OM_2-27-2026_12-05-56PM.OMN | S65 | 110 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.107 | 2.1 | 48.3 | TN | ug N/L | 1.23 | 15.7 | 916 |  | 1 | 1 |  | 48.3 | 48.3 | 48.3 | [6 empty] | 1 |  | 916 | 916 | 916 | [6 empty] | 148
2026-02-27 | 15:03 | OM_2-27-2026_12-05-56PM.OMN | DUP S65 | 116 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.0996 | 2.03 | 46.3 | TN | ug N/L | 1.04 | 13.2 | 756 |  | 1 | 1 |  | 46.3 | 46.3 | 46.3 |  |  | 15.7214 | PASS |  |  | 1 |  | 756 | 756 | 756 |  |  | 29.0909 | FAIL |  |  | 149
2026-02-27 | 15:04 | OM_2-27-2026_12-05-56PM.OMN | S66 | 111 | Unknown | 1 | 1 | TP | ug P/L | 0.114 | 2.31 | 54.2 | TN | ug N/L | 0.797 | 10.2 | 564 |  | 1 | 1 |  | 54.2 | 54.2 | 54.2 | [6 empty] | 1 |  | 564 | 564 | 564 | [6 empty] | 150
2026-02-27 | 15:05 | OM_2-27-2026_12-05-56PM.OMN | S67 | 112 | Unknown | 1 | 1 | TP | ug P/L | 0.153 | 2.97 | 73.6 | TN | ug N/L | 1.37 | 17.4 | 1020 |  | 1 | 1 |  | 73.6 | 73.6 | 73.6 | [6 empty] | 1 |  | 1020 | 1020 | 1020 | [6 empty] | 151
2026-02-27 | 15:06 | OM_2-27-2026_12-05-56PM.OMN | S68 | 113 | Unknown | 1 | 1 | TP | ug P/L | 0.0676 | 1.42 | 29.4 | TN | ug N/L | 0.68 | 8.69 | 469 |  | 1 | 1 |  | 29.4 | 29.4 | 29.4 | [6 empty] | 1 |  | 469 | 469 | 469 | [6 empty] | 152
2026-02-27 | 15:08 | OM_2-27-2026_12-05-56PM.OMN | S69 | 114 | Unknown | 1 | 1 | TP | ug P/L | 0.172 | 3.36 | 85.1 | TN | ug N/L | 1.11 | 14 | 808 |  | 1 | 1 |  | 85.1 | 85.1 | 85.1 | [6 empty] | 1 |  | 808 | 808 | 808 | [6 empty] | 153
2026-02-27 | 15:09 | OM_2-27-2026_12-05-56PM.OMN | S70 | 115 | Spike Set | 1 | 1 | TP | ug P/L | 0.0866 | 1.82 | 40.3 | TN | ug N/L | 0.768 | 9.87 | 543 |  | 1 | 1 |  | 40.3 | 40.3 | 40.3 | [6 empty] | 1 |  | 543 | 543 | 543 | [6 empty] | 154
2026-02-27 | 15:10 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S70 | 117 | Spike Set | 1 | 1 | TP | ug P/L | 0.14 | 2.76 | 67.2 | TN | ug N/L | 1.13 | 14.5 | 836 |  | 1 | 1 |  | 67.2 | 67.2 | 67.2 |  |  |  |  | 114.32 | PASS | 1 |  | 836 | 836 | 836 |  |  |  |  | 125.56 | FAIL | 155
2026-02-27 | 15:12 | OM_2-27-2026_12-05-56PM.OMN | CHECK 150 | 2 | Check Standard | 1 | 1 | TP | ug P/L | 0.291 | 5.47 | 153 | TN | ug N/L | 1.91 | 24.2 | 1480 |  | 1 | 1 |  | 153 | 153 | 153 | 2 | PASS |  |  |  |  | 1 |  | 1480 | 1480 | 1480 | -1.3333 | PASS |  |  |  |  | 156
2026-02-27 | 15:13 | OM_2-27-2026_12-05-56PM.OMN | CHECK 100 | 4 | Check Standard | 1 | 1 | TP | ug P/L | 0.223 | 3.76 | 97.4 | TN | ug N/L | 1.35 | 17.1 | 1010 |  | 1 | 1 |  | 97.4 | 97.4 | 97.4 | -2.6 | PASS |  |  |  |  | 1 |  | 1010 | 1010 | 1010 | 1 | PASS |  |  |  |  | 157
2026-02-27 | 15:14 | OM_2-27-2026_12-05-56PM.OMN | CHECK 50 | 6 | Check Standard | 1 | 1 | TP | ug P/L | 0.0967 | 1.97 | 44.6 | TN | ug N/L | 0.709 | 9.11 | 495 |  | 1 | 1 |  | 44.6 | 44.6 | 44.6 | -10.8 | PASS |  |  |  |  | 1 |  | 495 | 495 | 495 | -1 | PASS |  |  |  |  | 158
2026-02-27 | 15:15 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 | 8 | Check Standard | 1 | 1 | TP | ug P/L | 0.0515 | 1.13 | 21.5 | TN | ug N/L | 0.383 | 4.99 | 238 |  | 1 | 1 |  | 21.5 | 21.5 | 21.5 | -14 | PASS |  |  |  |  | 1 |  | 238 | 238 | 238 | -4.8 | PASS |  |  |  |  | 159
2026-02-27 | 15:17 | OM_2-27-2026_12-05-56PM.OMN | CHECK 10 | 10 | Check Standard | 1 | 1 | TP | ug P/L | 0.0255 | 0.717 | 10.4 | TN | ug N/L | 0.192 | 2.58 | 89.8 |  | 1 | 1 |  | 10.4 | 10.4 | 10.4 | 4 | PASS |  |  |  |  | 1 |  | 89.8 | 89.8 | 89.8 | -10.2 | PASS |  |  |  |  | 160
2026-02-27 | 15:18 | OM_2-27-2026_12-05-56PM.OMN | CHECK 5 | 12 | Check Standard | 1 | 1 | TP | ug P/L | 0.0147 | 0.42 | 2.7 | TN | ug N/L | 0.132 | 1.86 | 45.4 |  | 1 | 1 |  | 2.7 | 2.7 | 2.7 | -46 | FAIL |  |  |  |  | 1 |  | 45.4 | 45.4 | 45.4 | -9.2 | PASS |  |  |  |  | 161
2026-02-27 | 15:19 | OM_2-27-2026_12-05-56PM.OMN | CHECK 0 | 14 | Check Standard | 1 | 1 | TP | ug P/L | 0.0115 | 0.25 | -1.66 | TN | ug N/L | 0.0703 | 1.06 | -2.71 |  | 1 | 1 |  | -1.66 | -1.66 | -1.66 | -106.64 | FAIL |  |  |  |  | 1 |  | -2.71 | -2.71 | -2.71 | -101.084 | FAIL |  |  |  |  | 162
2026-02-27 | 15:20 | OM_2-27-2026_12-05-56PM.OMN | S71 | 118 | Unknown | 1 | 1 | TP | ug P/L | 0.243 | 4.65 | 125 | TN | ug N/L | 3.12 | 39.9 | 2570 |  | 1 | 1 |  | 125 | 125 | 125 | [6 empty] | 3 | out of range | 2570 | 2570 | 2570 | [6 empty] | 163
2026-02-27 | 15:22 | OM_2-27-2026_12-05-56PM.OMN | S72 | 119 | Unknown | 1 | 1 | TP | ug P/L | 0.118 | 2.42 | 57.4 | TN | ug N/L | 0.474 | 6.15 | 310 |  | 1 | 1 |  | 57.4 | 57.4 | 57.4 | [6 empty] | 1 |  | 310 | 310 | 310 | [6 empty] | 164
2026-02-27 | 15:23 | OM_2-27-2026_12-05-56PM.OMN | S73 | 120 | Unknown | 1 | 1 | TP | ug P/L | 0.204 | 3.91 | 102 | TN | ug N/L | 0.301 | 3.91 | 171 |  | 1 | 1 |  | 102 | 102 | 102 | [6 empty] | 1 |  | 171 | 171 | 171 | [6 empty] | 165
2026-02-27 | 15:24 | OM_2-27-2026_12-05-56PM.OMN | S74 | 121 | Unknown | 1 | 1 | TP | ug P/L | 0.828 | 15.3 | 556 | TN | ug N/L | 3.24 | 41.2 | 2660 |  | 1 | 3 | out of range | 556 | 556 | 556 | [6 empty] | 3 | out of range | 2660 | 2660 | 2660 | [6 empty] | 166
2026-02-27 | 15:26 | OM_2-27-2026_12-05-56PM.OMN | S75 | 122 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.286 | 5.5 | 154 | TN | ug N/L | 0.432 | 5.56 | 273 |  | 1 | 1 |  | 154 | 154 | 154 | [6 empty] | 1 |  | 273 | 273 | 273 | [6 empty] | 167
2026-02-27 | 15:27 | OM_2-27-2026_12-05-56PM.OMN | DUP S75 | 128 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.26 | 4.95 | 136 | TN | ug N/L | 0.4 | 5.16 | 248 |  | 1 | 1 |  | 136 | 136 | 136 |  |  | 12.4138 | PASS |  |  | 1 |  | 248 | 248 | 248 |  |  | 9.5969 | PASS |  |  | 168
2026-02-27 | 15:28 | OM_2-27-2026_12-05-56PM.OMN | S76 | 123 | Unknown | 1 | 1 | TP | ug P/L | 0.186 | 3.67 | 94.6 | TN | ug N/L | 1.49 | 18.9 | 1130 |  | 1 | 1 |  | 94.6 | 94.6 | 94.6 | [6 empty] | 1 |  | 1130 | 1130 | 1130 | [6 empty] | 169
2026-02-27 | 15:29 | OM_2-27-2026_12-05-56PM.OMN | S77 | 124 | Unknown | 1 | 1 | TP | ug P/L | 0.125 | 2.52 | 60.1 | TN | ug N/L | 0.227 | 3.02 | 116 |  | 1 | 1 |  | 60.1 | 60.1 | 60.1 | [6 empty] | 1 |  | 116 | 116 | 116 | [6 empty] | 170
2026-02-27 | 15:31 | OM_2-27-2026_12-05-56PM.OMN | S78 | 125 | Unknown | 1 | 1 | TP | ug P/L | 0.097 | 1.62 | 34.8 | TN | ug N/L | 0.928 | 11.8 | 667 |  | 1 | 1 |  | 34.8 | 34.8 | 34.8 | [6 empty] | 1 |  | 667 | 667 | 667 | [6 empty] | 171
2026-02-27 | 15:32 | OM_2-27-2026_12-05-56PM.OMN | S79 | 126 | Unknown | 1 | 1 | TP | ug P/L | 0.402 | 7.67 | 230 | TN | ug N/L | 1.37 | 17.5 | 1030 |  | 1 | 3 | out of range | 230 | 230 | 230 | [6 empty] | 1 |  | 1030 | 1030 | 1030 | [6 empty] | 172
2026-02-27 | 15:33 | OM_2-27-2026_12-05-56PM.OMN | S80 | 127 | Spike Set | 1 | 1 | TP | ug P/L | 0.314 | 5.94 | 169 | TN | ug N/L | 0.281 | 3.8 | 164 |  | 1 | 1 |  | 169 | 169 | 169 | [6 empty] | 1 |  | 164 | 164 | 164 | [6 empty] | 173
2026-02-27 | 15:35 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S80 | 129 | Spike Set | 1 | 1 | TP | ug P/L | 0.433 | 6.88 | 202 | TN | ug N/L | 0.555 | 6.65 | 340 |  | 1 | 3 | out of range | 202 | 202 | 202 |  |  |  |  | 152.2 | FAIL | 1 |  | 340 | 340 | 340 |  |  |  |  | 73.8 | FAIL | 174
2026-02-27 | 15:36 | OM_2-27-2026_12-05-56PM.OMN | BISULFATE | S4 | Unknown | 1 | 1 | TP | ug P/L | 0.00889 | 0.0925 | -5.66 | TN | ug N/L | 0.0187 | 0.376 | -44.4 |  | 1 | 1 |  | -5.66 | -5.66 | -5.66 | [6 empty] | 1 |  | -44.4 | -44.4 | -44.4 | [6 empty] | 175
2026-02-27 | 15:37 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 | 17 | Check Standard | 1 | 1 | TP | ug P/L | 0.0534 | 1.12 | 21.1 | TN | ug N/L | 0.376 | 4.9 | 232 |  | 1 | 1 |  | 21.1 | 21.1 | 21.1 | -15.6 | PASS |  |  |  |  | 1 |  | 232 | 232 | 232 | -7.2 | PASS |  |  |  |  | 176
2026-02-27 | 15:38 | OM_2-27-2026_12-05-56PM.OMN | NO2 | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.158 | 2.12 | 48.8 | TN | ug N/L | 1.86 | 24.3 | 1480 | 86.0082 | 1 | 1 |  | 48.8 | 48.8 | 48.8 | [6 empty] | 1 |  | 1480 | 1480 | 1480 | [6 empty] | 177
2026-02-27 | 15:40 | OM_2-27-2026_12-05-56PM.OMN | NO3 | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.142 | 1.97 | 44.5 | TN | ug N/L | 1.64 | 20.9 | 1260 |  | 1 | 1 |  | 44.5 | 44.5 | 44.5 | [6 empty] | 1 |  | 1260 | 1260 | 1260 | [6 empty] | 178
2026-02-27 | 15:41 | OM_2-27-2026_12-05-56PM.OMN | S81 | 130 | Unknown | 1 | 1 | TP | ug P/L | 0.0783 | 1.74 | 38.1 | TN | ug N/L | 0.824 | 10.4 | 576 |  | 1 | 1 |  | 38.1 | 38.1 | 38.1 | [6 empty] | 1 |  | 576 | 576 | 576 | [6 empty] | 179
2026-02-27 | 15:42 | OM_2-27-2026_12-05-56PM.OMN | S82 | 131 | Unknown | 1 | 1 | TP | ug P/L | 0.0991 | 1.99 | 45.1 | TN | ug N/L | 1.15 | 14.7 | 851 |  | 1 | 1 |  | 45.1 | 45.1 | 45.1 | [6 empty] | 1 |  | 851 | 851 | 851 | [6 empty] | 180
2026-02-27 | 15:44 | OM_2-27-2026_12-05-56PM.OMN | S83 | 132 | Unknown | 1 | 1 | TP | ug P/L | 0.19 | 3.6 | 92.5 | TN | ug N/L | 1.89 | 23.9 | 1460 |  | 1 | 1 |  | 92.5 | 92.5 | 92.5 | [6 empty] | 1 |  | 1460 | 1460 | 1460 | [6 empty] | 181
2026-02-27 | 15:45 | OM_2-27-2026_12-05-56PM.OMN | S84 | 133 | Unknown | 1 | 1 | TP | ug P/L | 0.0565 | 1.22 | 23.9 | TN | ug N/L | 0.658 | 8.45 | 453 |  | 1 | 1 |  | 23.9 | 23.9 | 23.9 | [6 empty] | 1 |  | 453 | 453 | 453 | [6 empty] | 182
2026-02-27 | 15:46 | OM_2-27-2026_12-05-56PM.OMN | S85 | 134 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.074 | 1.64 | 35.2 | TN | ug N/L | 0.675 | 8.67 | 467 |  | 1 | 2 | slightly messy peak | 35.2 | 35.2 | 35.2 | [6 empty] | 1 |  | 467 | 467 | 467 | [6 empty] | 183
2026-02-27 | 15:47 | OM_2-27-2026_12-05-56PM.OMN | DUP S85 | 140 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.0701 | 1.54 | 32.5 | TN | ug N/L | 0.697 | 8.97 | 486 |  | 1 | 2 | slightly messy peak | 32.5 | 32.5 | 32.5 |  |  | 7.9764 | PASS |  |  | 1 |  | 486 | 486 | 486 |  |  | 3.9874 | PASS |  |  | 184
2026-02-27 | 15:49 | OM_2-27-2026_12-05-56PM.OMN | S86 | 135 | Unknown | 1 | 1 | TP | ug P/L | 0.0659 | 1.68 | 36.4 | TN | ug N/L | 0.792 | 10.1 | 560 |  | 1 | 1 |  | 36.4 | 36.4 | 36.4 | [6 empty] | 1 |  | 560 | 560 | 560 | [6 empty] | 185
2026-02-27 | 15:50 | OM_2-27-2026_12-05-56PM.OMN | S87 | 136 | Unknown | 1 | 1 | TP | ug P/L | 0.739 | 13.4 | 468 | TN | ug N/L | 3.94 | 50.9 | 3380 |  | 1 | 3 | out of range | 468 | 468 | 468 | [6 empty] | 3 | out of range | 3380 | 3380 | 3380 | [6 empty] | 186
2026-02-27 | 15:51 | OM_2-27-2026_12-05-56PM.OMN | S88 | 137 | Unknown | 1 | 1 | TP | ug P/L | 0.133 | 2.72 | 66 | TN | ug N/L | 0.628 | 8.11 | 432 |  | 1 | 1 |  | 66 | 66 | 66 | [6 empty] | 1 |  | 432 | 432 | 432 | [6 empty] | 187
2026-02-27 | 15:53 | OM_2-27-2026_12-05-56PM.OMN | S89 | 138 | Unknown | 1 | 1 | TP | ug P/L | 0.234 | 4.46 | 120 | TN | ug N/L | 2.4 | 30.4 | 1900 |  | 1 | 1 |  | 120 | 120 | 120 | [6 empty] | 3 | out of range | 1900 | 1900 | 1900 | [6 empty] | 188
2026-02-27 | 15:54 | OM_2-27-2026_12-05-56PM.OMN | S90 | 139 | Spike Set | 1 | 1 | TP | ug P/L | 0.104 | 2.13 | 49 | TN | ug N/L | 0.233 | 3.13 | 123 |  | 1 | 1 |  | 49 | 49 | 49 | [6 empty] | 1 |  | 123 | 123 | 123 | [6 empty] | 189
2026-02-27 | 15:55 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S90 | 141 | Spike Set | 1 | 1 | TP | ug P/L | 0.156 | 3.03 | 75.3 | TN | ug N/L | 0.538 | 6.96 | 360 |  | 1 | 1 |  | 75.3 | 75.3 | 75.3 |  |  |  |  | 112.73 | PASS | 1 |  | 360 | 360 | 360 |  |  |  |  | 98.4 | PASS | 190
2026-02-27 | 15:56 | OM_2-27-2026_12-05-56PM.OMN | BLANK | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.138 | 1.67 | 36.1 | TN | ug N/L | 0.0131 | 0.293 | -49.4 |  | 1 | 1 |  | 36.1 | 36.1 | 36.1 | [6 empty] | 1 |  | -49.4 | -49.4 | -49.4 | [6 empty] | 191
2026-02-27 | 15:58 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 | 18 | Check Standard | 1 | 1 | TP | ug P/L | 0.0513 | 1.1 | 20.5 | TN | ug N/L | 0.405 | 5.16 | 248 |  | 1 | 1 |  | 20.5 | 20.5 | 20.5 | -18 | PASS |  |  |  |  | 1 |  | 248 | 248 | 248 | -0.8 | PASS |  |  |  |  | 192
2026-02-27 | 15:59 | OM_2-27-2026_12-05-56PM.OMN | S91 | 142 | Unknown | 1 | 1 | TP | ug P/L | 0.585 | 9.06 | 283 | TN | ug N/L | 5.02 | 65.1 | 4470 |  | 1 | 3 | out of range | 283 | 283 | 283 | [6 empty] | 3 | out of range | 4470 | 4470 | 4470 | [6 empty] | 193
2026-02-27 | 16:00 | OM_2-27-2026_12-05-56PM.OMN | S92 | 143 | Unknown | 1 | 1 | TP | ug P/L | 0.31 | 5.68 | 160 | TN | ug N/L | 0.422 | 5.44 | 266 |  | 1 | 1 |  | 160 | 160 | 160 | [6 empty] | 1 |  | 266 | 266 | 266 | [6 empty] | 194
2026-02-27 | 16:02 | OM_2-27-2026_12-05-56PM.OMN | S93 | 144 | Unknown | 1 | 1 | TP | ug P/L | 0.144 | 2.73 | 66.3 | TN | ug N/L | 0.94 | 12.2 | 688 |  | 1 | 1 |  | 66.3 | 66.3 | 66.3 | [6 empty] | 1 |  | 688 | 688 | 688 | [6 empty] | 195
2026-02-27 | 16:03 | OM_2-27-2026_12-05-56PM.OMN | S94 | 145 | Unknown | 1 | 1 | TP | ug P/L | 0.267 | 5.02 | 138 | TN | ug N/L | 1.6 | 20.3 | 1220 |  | 1 | 1 |  | 138 | 138 | 138 | [6 empty] | 1 |  | 1220 | 1220 | 1220 | [6 empty] | 196
2026-02-27 | 16:04 | OM_2-27-2026_12-05-56PM.OMN | S95 | 146 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.392 | 7.24 | 215 | TN | ug N/L | 1.48 | 18.8 | 1120 |  | 1 | 3 | out of range | 215 | 215 | 215 | [6 empty] | 1 |  | 1120 | 1120 | 1120 | [6 empty] | 197
2026-02-27 | 16:05 | OM_2-27-2026_12-05-56PM.OMN | DUP S95 | 152 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.397 | 7.3 | 217 | TN | ug N/L | 1.56 | 19.8 | 1190 |  | 1 | 3 | out of range | 217 | 217 | 217 |  |  | 0.92593 | PASS |  |  | 1 |  | 1190 | 1190 | 1190 |  |  | 6.0606 | PASS |  |  | 198
2026-02-27 | 16:07 | OM_2-27-2026_12-05-56PM.OMN | S96 | 147 | Unknown | 1 | 1 | TP | ug P/L | 0.25 | 4.43 | 119 | TN | ug N/L | 1.72 | 21.8 | 1320 |  | 1 | 1 |  | 119 | 119 | 119 | [6 empty] | 1 |  | 1320 | 1320 | 1320 | [6 empty] | 199
2026-02-27 | 16:08 | OM_2-27-2026_12-05-56PM.OMN | S97 | 148 | Unknown | 1 | 1 | TP | ug P/L | 0.215 | 4.16 | 110 | TN | ug N/L | 1.44 | 18.3 | 1080 |  | 1 | 1 |  | 110 | 110 | 110 | [6 empty] | 1 |  | 1080 | 1080 | 1080 | [6 empty] | 200
2026-02-27 | 16:09 | OM_2-27-2026_12-05-56PM.OMN | S98 | 149 | Unknown | 1 | 1 | TP | ug P/L | 0.115 | 2.14 | 49.3 | TN | ug N/L | 1.39 | 17.6 | 1040 |  | 1 | 1 |  | 49.3 | 49.3 | 49.3 | [6 empty] | 1 |  | 1040 | 1040 | 1040 | [6 empty] | 201
2026-02-27 | 16:11 | OM_2-27-2026_12-05-56PM.OMN | S99 | 150 | Unknown | 1 | 1 | TP | ug P/L | 0.0873 | 1.8 | 39.7 | TN | ug N/L | 0.724 | 9.29 | 506 |  | 1 | 1 |  | 39.7 | 39.7 | 39.7 | [6 empty] | 1 |  | 506 | 506 | 506 | [6 empty] | 202
2026-02-27 | 16:12 | OM_2-27-2026_12-05-56PM.OMN | S100 | 151 | Spike Set | 1 | 1 | TP | ug P/L | 0.0744 | 1.63 | 35 | TN | ug N/L | 0.77 | 9.93 | 547 |  | 1 | 1 |  | 35 | 35 | 35 | [6 empty] | 1 |  | 547 | 547 | 547 | [6 empty] | 203
2026-02-27 | 16:13 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S100 | 153 | Spike Set | 1 | 1 | TP | ug P/L | 0.123 | 2.54 | 60.9 | TN | ug N/L | 1.03 | 13.2 | 756 |  | 1 | 1 |  | 60.9 | 60.9 | 60.9 |  |  |  |  | 109.69 | PASS | 1 |  | 756 | 756 | 756 |  |  |  |  | 91.16 | PASS | 204
2026-02-27 | 16:14 | OM_2-27-2026_12-05-56PM.OMN | BISULFATE | S4 | Unknown | 1 | 1 | TP | ug P/L | 0.00999 | 0.107 | -5.29 | TN | ug N/L | 0.0191 | 0.317 | -48 |  | 1 | 1 |  | -5.29 | -5.29 | -5.29 | [6 empty] | 1 |  | -48 | -48 | -48 | [6 empty] | 205
2026-02-27 | 16:16 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 | 15 | Check Standard | 1 | 1 | TP | ug P/L | 0.0503 | 1.17 | 22.4 | TN | ug N/L | 0.384 | 5.05 | 241 |  | 1 | 1 |  | 22.4 | 22.4 | 22.4 | -10.4 | PASS |  |  |  |  | 1 |  | 241 | 241 | 241 | -3.6 | PASS |  |  |  |  | 206
2026-02-27 | 16:17 | OM_2-27-2026_12-05-56PM.OMN | NO2 | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.139 | 1.73 | 37.9 | TN | ug N/L | 1.86 | 24.4 | 1490 | 113.1148 | 1 | 1 |  | 37.9 | 37.9 | 37.9 | [6 empty] | 1 |  | 1490 | 1490 | 1490 | [6 empty] | 207
2026-02-27 | 16:18 | OM_2-27-2026_12-05-56PM.OMN | NO3 | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.147 | 1.87 | 41.7 | TN | ug N/L | 2.17 | 27.6 | 1710 |  | 1 | 1 |  | 41.7 | 41.7 | 41.7 | [6 empty] | 1 |  | 1710 | 1710 | 1710 | [6 empty] | 208
2026-02-27 | 16:20 | OM_2-27-2026_12-05-56PM.OMN | S101 | 154 | Unknown | 1 | 1 | TP | ug P/L | 0.0663 | 1.42 | 29.3 | TN | ug N/L | 0.516 | 6.71 | 344 |  | 1 | 1 |  | 29.3 | 29.3 | 29.3 | [6 empty] | 1 |  | 344 | 344 | 344 | [6 empty] | 209
2026-02-27 | 16:21 | OM_2-27-2026_12-05-56PM.OMN | S102 | 155 | Unknown | 1 | 1 | TP | ug P/L | 0.0914 | 1.94 | 43.6 | TN | ug N/L | 0.448 | 5.86 | 292 |  | 1 | 1 |  | 43.6 | 43.6 | 43.6 | [6 empty] | 1 |  | 292 | 292 | 292 | [6 empty] | 210
2026-02-27 | 16:22 | OM_2-27-2026_12-05-56PM.OMN | S103 | 156 | Unknown | 1 | 1 | TP | ug P/L | 0.0866 | 1.84 | 41 | TN | ug N/L | 0.645 | 8.33 | 446 |  | 1 | 1 |  | 41 | 41 | 41 | [6 empty] | 1 |  | 446 | 446 | 446 | [6 empty] | 211
2026-02-27 | 16:23 | OM_2-27-2026_12-05-56PM.OMN | S104 | 157 | Unknown | 1 | 1 | TP | ug P/L | 0.101 | 2.1 | 48.2 | TN | ug N/L | 1.33 | 16.8 | 990 |  | 1 | 1 |  | 48.2 | 48.2 | 48.2 | [6 empty] | 1 |  | 990 | 990 | 990 | [6 empty] | 212
2026-02-27 | 16:25 | OM_2-27-2026_12-05-56PM.OMN | S105 | 158 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.064 | 1.38 | 28.2 | TN | ug N/L | 0.68 | 8.86 | 479 |  | 1 | 1 |  | 28.2 | 28.2 | 28.2 | [6 empty] | 1 |  | 479 | 479 | 479 | [6 empty] | 213
2026-02-27 | 16:26 | OM_2-27-2026_12-05-56PM.OMN | DUP S105 | 164 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.0876 | 1.38 | 28.3 | TN | ug N/L | 0.79 | 8.81 | 476 |  | 1 | 1 |  | 28.3 | 28.3 | 28.3 |  |  | 0.35398 | PASS |  |  | 1 |  | 476 | 476 | 476 |  |  | 0.62827 | PASS |  |  | 214
2026-02-27 | 16:27 | OM_2-27-2026_12-05-56PM.OMN | S106 | 159 | Unknown | 1 | 1 | TP | ug P/L | 0.128 | 2.59 | 62.2 | TN | ug N/L | 0.827 | 9.63 | 527 |  | 1 | 1 |  | 62.2 | 62.2 | 62.2 | [6 empty] | 1 |  | 527 | 527 | 527 | [6 empty] | 215
2026-02-27 | 16:28 | OM_2-27-2026_12-05-56PM.OMN | S107 | 160 | Unknown | 1 | 1 | TP | ug P/L | 0.0624 | 1.39 | 28.6 | TN | ug N/L | 0.608 | 7.91 | 419 |  | 1 | 1 |  | 28.6 | 28.6 | 28.6 | [6 empty] | 1 |  | 419 | 419 | 419 | [6 empty] | 216
2026-02-27 | 16:30 | OM_2-27-2026_12-05-56PM.OMN | S108 | 161 | Unknown | 1 | 1 | TP | ug P/L | 0.0848 | 1.79 | 39.3 | TN | ug N/L | 0.981 | 12.6 | 715 |  | 1 | 1 |  | 39.3 | 39.3 | 39.3 | [6 empty] | 1 |  | 715 | 715 | 715 | [6 empty] | 217
2026-02-27 | 16:31 | OM_2-27-2026_12-05-56PM.OMN | S109 | 162 | Unknown | 1 | 1 | TP | ug P/L | 0.0556 | 1.29 | 25.8 | TN | ug N/L | 0.604 | 7.81 | 413 |  | 1 | 1 |  | 25.8 | 25.8 | 25.8 | [6 empty] | 1 |  | 413 | 413 | 413 | [6 empty] | 218
2026-02-27 | 16:32 | OM_2-27-2026_12-05-56PM.OMN | S110 | 163 | Spike Set | 1 | 1 | TP | ug P/L | 0.242 | 4.63 | 125 | TN | ug N/L | 2.55 | 32.5 | 2050 |  | 1 | 1 |  | 125 | 125 | 125 | [6 empty] | 3 | out of range | 2050 | 2050 | 2050 | [6 empty] | 219
2026-02-27 | 16:34 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S110 | 165 | Spike Set | 1 | 1 | TP | ug P/L | 0.18 | 3.51 | 89.8 | TN | ug N/L | 1.6 | 20.4 | 1230 |  | 1 | 1 |  | 89.8 | 89.8 | 89.8 |  |  |  |  | -131.82 | FAIL | 1 |  | 1230 | 1230 | 1230 |  |  |  |  | -315.7 | FAIL | 220
2026-02-27 | 16:35 | OM_2-27-2026_12-05-56PM.OMN | BLANK | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.138 | 1.8 | 39.7 | TN | ug N/L | 0.00349 | 0.0185 | -66 |  | 1 | 1 |  | 39.7 | 39.7 | 39.7 | [6 empty] | 1 |  | -66 | -66 | -66 | [6 empty] | 221
2026-02-27 | 16:36 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 | 16 | Check Standard | 1 | 1 | TP | ug P/L | 0.0506 | 1.09 | 20.4 | TN | ug N/L | 0.354 | 4.71 | 221 |  | 1 | 1 |  | 20.4 | 20.4 | 20.4 | -18.4 | PASS |  |  |  |  | 1 |  | 221 | 221 | 221 | -11.6 | PASS |  |  |  |  | 222
2026-02-27 | 16:37 | OM_2-27-2026_12-05-56PM.OMN | S111 | 166 | Unknown | 1 | 1 | TP | ug P/L | 0.167 | 3.23 | 81.3 | TN | ug N/L | 0.389 | 5.07 | 243 |  | 1 | 1 |  | 81.3 | 81.3 | 81.3 | [6 empty] | 1 |  | 243 | 243 | 243 | [6 empty] | 223
2026-02-27 | 16:39 | OM_2-27-2026_12-05-56PM.OMN | S112 | 167 | Unknown | 1 | 1 | TP | ug P/L | 1.5 | 27.1 | 1250 | TN | ug N/L | 5.6 | 73.8 | 5170 |  | 1 | 3 | out of range | 1250 | 1250 | 1250 | [6 empty] | 3 | out of range | 5170 | 5170 | 5170 | [6 empty] | 224
2026-02-27 | 16:40 | OM_2-27-2026_12-05-56PM.OMN | S113 | 168 | Unknown | 1 | 1 | TP | ug P/L | 0.165 | 2.85 | 69.9 | TN | ug N/L | 1.39 | 17.8 | 1050 |  | 1 | 1 |  | 69.9 | 69.9 | 69.9 | [6 empty] | 1 |  | 1050 | 1050 | 1050 | [6 empty] | 225
2026-02-27 | 16:41 | OM_2-27-2026_12-05-56PM.OMN | S114 | 169 | Unknown | 1 | 1 | TP | ug P/L | 0.0764 | 1.59 | 33.9 | TN | ug N/L | 0.744 | 7.83 | 415 |  | 1 | 1 |  | 33.9 | 33.9 | 33.9 | [6 empty] | 1 |  | 415 | 415 | 415 | [6 empty] | 226
2026-02-27 | 16:43 | OM_2-27-2026_12-05-56PM.OMN | S115 | 170 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.0814 | 1.76 | 38.7 | TN | ug N/L | 0.827 | 10.6 | 592 |  | 1 | 3 | messy P- use dup | 38.7 | 38.7 | 38.7 | [6 empty] | 1 |  | 592 | 592 | 592 | [6 empty] | 227
2026-02-27 | 16:44 | OM_2-27-2026_12-05-56PM.OMN | DUP S115 | 176 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.0568 | 1.41 | 29.1 | TN | ug N/L | 0.662 | 8.58 | 462 |  | 1 | 2 | original had messy integration | 29.1 | 29.1 | 29.1 |  |  | 28.3186 | FAIL |  |  | 1 |  | 462 | 462 | 462 |  |  | 24.6679 | FAIL |  |  | 228
2026-02-27 | 16:45 | OM_2-27-2026_12-05-56PM.OMN | S116 | 171 | Unknown | 1 | 1 | TP | ug P/L | 0.12 | 2.44 | 58 | TN | ug N/L | 1.12 | 14.3 | 828 |  | 1 | 1 |  | 58 | 58 | 58 | [6 empty] | 1 |  | 828 | 828 | 828 | [6 empty] | 229
2026-02-27 | 16:46 | OM_2-27-2026_12-05-56PM.OMN | S117 | 172 | Unknown | 1 | 1 | TP | ug P/L | 0.862 | 15.7 | 576 | TN | ug N/L | 3.73 | 47.8 | 3150 |  | 1 | 3 | out of range | 576 | 576 | 576 | [6 empty] | 3 | out of range | 3150 | 3150 | 3150 | [6 empty] | 230
2026-02-27 | 16:48 | OM_2-27-2026_12-05-56PM.OMN | S118 | 173 | Unknown | 1 | 1 | TP | ug P/L | 0.414 | 7.75 | 234 | TN | ug N/L | 2.18 | 27.8 | 1720 |  | 1 | 3 | out of range | 234 | 234 | 234 | [6 empty] | 1 |  | 1720 | 1720 | 1720 | [6 empty] | 231
2026-02-27 | 16:49 | OM_2-27-2026_12-05-56PM.OMN | S119 | 174 | Unknown | 1 | 1 | TP | ug P/L | 0.0861 | 1.81 | 40 | TN | ug N/L | 0.768 | 9.92 | 546 |  | 1 | 1 |  | 40 | 40 | 40 | [6 empty] | 1 |  | 546 | 546 | 546 | [6 empty] | 232
2026-02-27 | 16:50 | OM_2-27-2026_12-05-56PM.OMN | S120 | 175 | Spike Set | 1 | 1 | TP | ug P/L | 0.244 | 4.72 | 128 | TN | ug N/L | 1.56 | 19.9 | 1190 |  | 1 | 1 |  | 128 | 128 | 128 | [6 empty] | 1 |  | 1190 | 1190 | 1190 | [6 empty] | 233
2026-02-27 | 16:52 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S120 | 177 | Spike Set | 1 | 1 | TP | ug P/L | 0.207 | 3.97 | 104 | TN | ug N/L | 1.35 | 17.2 | 1020 |  | 1 | 1 |  | 104 | 104 | 104 |  |  |  |  | -85.6 | FAIL | 1 |  | 1020 | 1020 | 1020 |  |  |  |  | -57.8 | FAIL | 234
2026-02-27 | 16:53 | OM_2-27-2026_12-05-56PM.OMN | BISULFATE | S4 | Unknown | 1 | 1 | TP | ug P/L | 0.00929 | 0.103 | -5.39 | TN | ug N/L | 0.0182 | 0.364 | -45.1 |  | 1 | 1 |  | -5.39 | -5.39 | -5.39 | [6 empty] | 1 |  | -45.1 | -45.1 | -45.1 | [6 empty] | 235
2026-02-27 | 16:54 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 | 17 | Check Standard | 1 | 1 | TP | ug P/L | 0.0507 | 1.15 | 21.9 | TN | ug N/L | 0.368 | 4.85 | 229 |  | 1 | 1 |  | 21.9 | 21.9 | 21.9 | -12.4 | PASS |  |  |  |  | 1 |  | 229 | 229 | 229 | -8.4 | PASS |  |  |  |  | 236
2026-02-27 | 16:55 | OM_2-27-2026_12-05-56PM.OMN | NO2 | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.137 | 1.66 | 35.8 | TN | ug N/L | 1.74 | 22.9 | 1390 | 106.5502 | 1 | 1 |  | 35.8 | 35.8 | 35.8 | [6 empty] | 1 |  | 1390 | 1390 | 1390 | [6 empty] | 237
2026-02-27 | 16:57 | OM_2-27-2026_12-05-56PM.OMN | NO3 | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.137 | 1.68 | 36.4 | TN | ug N/L | 1.9 | 24.4 | 1490 |  | 1 | 1 |  | 36.4 | 36.4 | 36.4 | [6 empty] | 1 |  | 1490 | 1490 | 1490 | [6 empty] | 238
2026-02-27 | 16:58 | OM_2-27-2026_12-05-56PM.OMN | S121 | 178 | Unknown | 1 | 1 | TP | ug P/L | 0.318 | 6.06 | 173 | TN | ug N/L | 0.991 | 12.8 | 732 |  | 1 | 1 |  | 173 | 173 | 173 | [6 empty] | 1 |  | 732 | 732 | 732 | [6 empty] | 239
2026-02-27 | 16:59 | OM_2-27-2026_12-05-56PM.OMN | S122 | 179 | Unknown | 1 | 1 | TP | ug P/L | 0.373 | 7.05 | 208 | TN | ug N/L | 1.36 | 15.5 | 905 |  | 1 | 3 | out of range | 208 | 208 | 208 | [6 empty] | 1 |  | 905 | 905 | 905 | [6 empty] | 240
2026-02-27 | 17:01 | OM_2-27-2026_12-05-56PM.OMN | S123 | 180 | Unknown | 1 | 1 | TP | ug P/L | 0.778 | 14.4 | 514 | TN | ug N/L | 3.09 | 39.2 | 2520 |  | 1 | 3 | out of range | 514 | 514 | 514 | [6 empty] | 3 | out of range | 2520 | 2520 | 2520 | [6 empty] | 241
2026-02-27 | 17:02 | OM_2-27-2026_12-05-56PM.OMN | S124 | 181 | Unknown | 1 | 1 | TP | ug P/L | 0.543 | 10.1 | 325 | TN | ug N/L | 1.69 | 21.4 | 1290 |  | 1 | 3 | out of range | 325 | 325 | 325 | [6 empty] | 1 |  | 1290 | 1290 | 1290 | [6 empty] | 242
2026-02-27 | 17:03 | OM_2-27-2026_12-05-56PM.OMN | S125 | 182 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.42 | 7.93 | 240 | TN | ug N/L | 3.21 | 40.8 | 2630 |  | 1 | 3 | out of range | 240 | 240 | 240 | [6 empty] | 3 | out of range | 2630 | 2630 | 2630 | [6 empty] | 243
2026-02-27 | 17:04 | OM_2-27-2026_12-05-56PM.OMN | DUP S125 | 188 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.254 | 4.89 | 133 | TN | ug N/L | 2.01 | 25.6 | 1570 |  | 1 | 1 |  | 133 | 133 | 133 |  |  | 57.3727 | FAIL |  |  | 1 |  | 1570 | 1570 | 1570 |  |  | 50.4762 | FAIL |  |  | 244
2026-02-27 | 17:06 | OM_2-27-2026_12-05-56PM.OMN | S126 | 183 | Unknown | 1 | 1 | TP | ug P/L | 0.406 | 7.25 | 215 | TN | ug N/L | 1.19 | 15.2 | 885 |  | 1 | 3 | out of range | 215 | 215 | 215 | [6 empty] | 1 |  | 885 | 885 | 885 | [6 empty] | 245
2026-02-27 | 17:07 | OM_2-27-2026_12-05-56PM.OMN | S127 | 184 | Unknown | 1 | 1 | TP | ug P/L | 0.889 | 16.4 | 614 | TN | ug N/L | 4.3 | 55.1 | 3690 |  | 1 | 3 | out of range | 614 | 614 | 614 | [6 empty] | 3 | out of range | 3690 | 3690 | 3690 | [6 empty] | 246
2026-02-27 | 17:08 | OM_2-27-2026_12-05-56PM.OMN | S128 | 185 | Unknown | 1 | 1 | TP | ug P/L | 0.316 | 5.99 | 170 | TN | ug N/L | 1.1 | 14 | 806 |  | 1 | 1 |  | 170 | 170 | 170 | [6 empty] | 1 |  | 806 | 806 | 806 | [6 empty] | 247
2026-02-27 | 17:10 | OM_2-27-2026_12-05-56PM.OMN | S129 | 186 | Unknown | 1 | 1 | TP | ug P/L | 0.72 | 13.3 | 464 | TN | ug N/L | 0.955 | 12.2 | 689 |  | 1 | 3 | out of range | 464 | 464 | 464 | [6 empty] | 1 |  | 689 | 689 | 689 | [6 empty] | 248
2026-02-27 | 17:11 | OM_2-27-2026_12-05-56PM.OMN | S130 | 187 | Spike Set | 1 | 1 | TP | ug P/L | 0.852 | 15.7 | 579 | TN | ug N/L | 4.66 | 60.3 | 4090 |  | 1 | 3 | out of range | 579 | 579 | 579 | [6 empty] | 3 | out of range | 4090 | 4090 | 4090 | [6 empty] | 249
2026-02-27 | 17:12 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S130 | 189 | Spike Set | 1 | 1 | TP | ug P/L | 0.754 | 13.9 | 493 | TN | ug N/L | 3.89 | 50.1 | 3320 |  | 1 | 3 | out of range | 493 | 493 | 493 |  |  |  |  | -294.7 | FAIL | 3 | out of range | 3320 | 3320 | 3320 |  |  |  |  | -274.8 | FAIL | 250
2026-02-27 | 17:14 | OM_2-27-2026_12-05-56PM.OMN | BLANK | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.136 | 1.69 | 36.7 | TN | ug N/L | 0.00781 | 0.181 | -56.2 |  | 1 | 1 |  | 36.7 | 36.7 | 36.7 | [6 empty] | 1 |  | -56.2 | -56.2 | -56.2 | [6 empty] | 251
2026-02-27 | 17:15 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 | 18 | Check Standard | 1 | 1 | TP | ug P/L | 0.0504 | 1.13 | 21.4 | TN | ug N/L | 0.339 | 4.47 | 206 |  | 1 | 1 |  | 21.4 | 21.4 | 21.4 | -14.4 | PASS |  |  |  |  | 1 |  | 206 | 206 | 206 | -17.6 | PASS |  |  |  |  | 252
2026-02-27 | 17:16 | OM_2-27-2026_12-05-56PM.OMN | S131 | 190 | Unknown | 1 | 1 | TP | ug P/L | 0.854 | 15.8 | 584 | TN | ug N/L | 2.99 | 38.1 | 2440 |  | 1 | 3 | out of range | 584 | 584 | 584 | [6 empty] | 3 | out of range | 2440 | 2440 | 2440 | [6 empty] | 253
2026-02-27 | 17:17 | OM_2-27-2026_12-05-56PM.OMN | S132 | 191 | Unknown | 1 | 1 | TP | ug P/L | 0.143 | 2.91 | 71.6 | TN | ug N/L | 0.666 | 8.6 | 463 |  | 1 | 1 |  | 71.6 | 71.6 | 71.6 | [6 empty] | 1 |  | 463 | 463 | 463 | [6 empty] | 254
2026-02-27 | 17:19 | OM_2-27-2026_12-05-56PM.OMN | S133 | 192 | Unknown | 1 | 1 | TP | ug P/L | 0.142 | 2.85 | 69.9 | TN | ug N/L | 1.11 | 14.2 | 819 |  | 1 | 1 |  | 69.9 | 69.9 | 69.9 | [6 empty] | 1 |  | 819 | 819 | 819 | [6 empty] | 255
2026-02-27 | 17:20 | OM_2-27-2026_12-05-56PM.OMN | S134 | 193 | Unknown | 1 | 1 | TP | ug P/L | 0.558 | 10.5 | 342 | TN | ug N/L | 1.56 | 19.8 | 1180 |  | 1 | 3 | out of range | 342 | 342 | 342 | [6 empty] | 1 |  | 1180 | 1180 | 1180 | [6 empty] | 256
2026-02-27 | 17:21 | OM_2-27-2026_12-05-56PM.OMN | S135 | 194 | Duplicate Set | 1 | 1 | TP | ug P/L | 1.13 | 20.5 | 834 | TN | ug N/L | 2.58 | 32.5 | 2050 |  | 1 | 3 | out of range | 834 | 834 | 834 | [6 empty] | 3 | out of range | 2050 | 2050 | 2050 | [6 empty] | 257
2026-02-27 | 17:23 | OM_2-27-2026_12-05-56PM.OMN | DUP S135 | 199 | Duplicate Set | 1 | 1 | TP | ug P/L | 1.4 | 25.5 | 1140 | TN | ug N/L | 3.43 | 43.4 | 2820 |  | 1 | 3 | out of range | 1140 | 1140 | 1140 |  |  | 31.003 | FAIL |  |  | 3 | out of range | 2820 | 2820 | 2820 |  |  | 31.6222 | FAIL |  |  | 258
2026-02-27 | 17:24 | OM_2-27-2026_12-05-56PM.OMN | S136 | 195 | Unknown | 1 | 1 | TP | ug P/L | 0.232 | 4.68 | 127 | TN | ug N/L | 0.316 | 4.1 | 183 |  | 1 | 1 |  | 127 | 127 | 127 | [6 empty] | 1 |  | 183 | 183 | 183 | [6 empty] | 259
2026-02-27 | 17:25 | OM_2-27-2026_12-05-56PM.OMN | S137 | 196 | Unknown | 1 | 1 | TP | ug P/L | 0.0791 | 1.61 | 34.6 | TN | ug N/L | 0.642 | 8.29 | 443 |  | 1 | 1 |  | 34.6 | 34.6 | 34.6 | [6 empty] | 1 |  | 443 | 443 | 443 | [6 empty] | 260
2026-02-27 | 17:26 | OM_2-27-2026_12-05-56PM.OMN | S138 | 197 | Unknown | 1 | 1 | TP | ug P/L | 0.134 | 2.66 | 64.2 | TN | ug N/L | 1.09 | 13.9 | 802 |  | 1 | 1 |  | 64.2 | 64.2 | 64.2 | [6 empty] | 1 |  | 802 | 802 | 802 | [6 empty] | 261
2026-02-27 | 17:28 | OM_2-27-2026_12-05-56PM.OMN | S139 | 198 | Spike Set | 1 | 1 | TP | ug P/L | 0.382 | 7.05 | 208 | TN | ug N/L | 1.17 | 15.1 | 874 |  | 1 | 3 | out of range | 208 | 208 | 208 | [6 empty] | 1 |  | 874 | 874 | 874 | [6 empty] | 262
2026-02-27 | 17:29 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S139 | 200 | Spike Set | 1 | 1 | TP | ug P/L | 0.443 | 8.08 | 246 | TN | ug N/L | 1.49 | 19.1 | 1140 |  | 1 | 3 | out of range | 246 | 246 | 246 |  |  |  |  | 176.6 | FAIL | 1 |  | 1140 | 1140 | 1140 |  |  |  |  | 117.8 | PASS | 263
2026-02-27 | 17:30 | OM_2-27-2026_12-05-56PM.OMN | CHECK 150 | 1 | Check Standard | 1 | 1 | TP | ug P/L | 0.3 | 5.46 | 152 | TN | ug N/L | 2.4 | 21.4 | 1290 |  | 1 | 1 |  | 152 | 152 | 152 | 1.3333 | PASS |  |  |  |  | 1 |  | 1290 | 1290 | 1290 | -14 | PASS |  |  |  |  | 264
2026-02-27 | 17:32 | OM_2-27-2026_12-05-56PM.OMN | CHECK 100 | 3 | Check Standard | 1 | 1 | TP | ug P/L | 0.197 | 3.69 | 95.3 | TN | ug N/L | 1.27 | 16.2 | 946 |  | 1 | 1 |  | 95.3 | 95.3 | 95.3 | -4.7 | PASS |  |  |  |  | 1 |  | 946 | 946 | 946 | -5.4 | PASS |  |  |  |  | 265
2026-02-27 | 17:33 | OM_2-27-2026_12-05-56PM.OMN | CHECK 50 | 5 | Check Standard | 1 | 1 | TP | ug P/L | 0.124 | 2.05 | 46.8 | TN | ug N/L | 0.652 | 8.4 | 450 |  | 1 | 1 |  | 46.8 | 46.8 | 46.8 | -6.4 | PASS |  |  |  |  | 1 |  | 450 | 450 | 450 | -10 | PASS |  |  |  |  | 266
2026-02-27 | 17:34 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 | 7 | Check Standard | 1 | 1 | TP | ug P/L | 0.0529 | 1.19 | 23.1 | TN | ug N/L | 0.397 | 5.19 | 250 |  | 1 | 1 |  | 23.1 | 23.1 | 23.1 | -7.6 | PASS |  |  |  |  | 1 |  | 250 | 250 | 250 | 0 | PASS |  |  |  |  | 267
2026-02-27 | 17:35 | OM_2-27-2026_12-05-56PM.OMN | CHECK 0 | 13 | Check Standard | 1 | 1 | TP | ug P/L | 0.00922 | 0.152 | -4.16 | TN | ug N/L | 0.0692 | 1.07 | -2.44 |  | 1 | 1 |  | -4.16 | -4.16 | -4.16 | [6 empty] | 1 |  | -2.44 | -2.44 | -2.44 | [6 empty] | 268
2026-02-27 | 17:37 | OM_2-27-2026_12-05-56PM.OMN | NO2 | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.137 | 1.69 | 36.6 | TN | ug N/L | 1.83 | 23.9 | 1460 | 99.1632 | 1 | 1 |  | 36.6 | 36.6 | 36.6 | [6 empty] | 1 |  | 1460 | 1460 | 1460 | [6 empty] | 269
2026-02-27 | 17:40 | OM_2-27-2026_12-05-56PM.OMN | NO3 | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.135 | 1.66 | 36 | TN | ug N/L | 1.86 | 23.7 | 1450 |  | 1 | 1 |  | 36 | 36 | 36 | [6 empty] | 1 |  | 1450 | 1450 | 1450 | [6 empty] | 270

## method checks
Detection Date | Detection Time | Run File Name | Sample ID | Cup Number | Sample Type | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Column efficiency | DILUTION BY OXIDATION REAGENT | Analyst QUALITY CODE PO4-P: | Analyst NOTES: | PO4-P-- determined concentration | PO4-P-- blank corrected | PO4-P-- dilution corrected | PO4-P- % Error in check stds | PO4-P- DQO (<=20%) | PO4-P- RPD of duplicates | PO4-P- DQO (<=20% if conc>LOQ) | PO4-P- PR of spikes | PO4-P- DQO (80-120%) if conc>LOQ | Analyst QUALITY CODE NO3-N: | Analyst NOTES: | NO3-N-- determined concentration | NO3-N-- blank corrected | NO3-N-- dilution corrected | NO3-N- % Error in check stds | NO3-N- DQO (<=20%) | NO3-N- RPD of duplicates | NO3-N- DQO (<=20% if conc>LOQ) | NO3-N- PR of spikes | NO3-N- DQO (80-120%) if conc>LOQ
BLANKS-digested type I reagent grade water
2026-02-27 | 12:42 | OM_2-27-2026_12-05-56PM.OMN | water blank | 19 | Unknown | 1 | 1 | TP | ug P/L | 0.0128 | 0.314 | -0.0151 | TN | ug N/L | 0.0703 | 1.1 | -0.261 |  | 1 | 1 |  | -0.0151 | -0.0151 | -0.0151 | [6 empty] | 1 |  | -0.261 | -0.261 | -0.261 | [6 empty] | 37
2026-02-27 | 12:43 | OM_2-27-2026_12-05-56PM.OMN | water blank | 20 | Unknown | 1 | 1 | TP | ug P/L | 0.0111 | 0.291 | -0.624 | TN | ug N/L | 0.0676 | 1.05 | -3.38 |  | 1 | 1 |  | -0.624 | -0.624 | -0.624 | [6 empty] | 1 |  | -3.38 | -3.38 | -3.38 | [6 empty] | 38
2026-02-27 | 12:45 | OM_2-27-2026_12-05-56PM.OMN | water blank | 21 | Unknown | 1 | 1 | TP | ug P/L | 0.0122 | 0.313 | -0.0603 | TN | ug N/L | 0.0688 | 1.06 | -2.77 |  | 1 | 1 |  | -0.0603 | -0.0603 | -0.0603 | [6 empty] | 1 |  | -2.77 | -2.77 | -2.77 | [6 empty] | 39
2026-02-27 | 12:46 | OM_2-27-2026_12-05-56PM.OMN | water blank | 22 | Unknown | 1 | 1 | TP | ug P/L | 0.0126 | 0.287 | -0.706 | TN | ug N/L | 0.069 | 1.06 | -2.88 |  | 1 | 1 |  | -0.706 | -0.706 | -0.706 | [6 empty] | 1 |  | -2.88 | -2.88 | -2.88 | [6 empty] | 40
2026-02-27 | 12:47 | OM_2-27-2026_12-05-56PM.OMN | water blank | 23 | Unknown | 1 | 1 | TP | ug P/L | 0.0124 | 0.312 | -0.0704 | TN | ug N/L | 0.0692 | 1.07 | -2.32 |  | 1 | 1 |  | -0.0704 | -0.0704 | -0.0704 | [6 empty] | 1 |  | -2.32 | -2.32 | -2.32 | [6 empty] | 41
[9 empty] | Known Conc |  |  | 0 |  | Known Conc |  |  | 0
[9 empty] | Mean |  | 0.3034 | -0.29516 |  | Mean |  | 1.068 | -2.3222
[9 empty] | Std Dev |  | 0.01324 | 0.3395 |  | Std Dev |  | 0.019235 | 1.2123
[9 empty] | CV or %RSD |  | 4.3639 | -115.0216 |  | CV or %RSD |  | 1.8011 | -52.2069
[9 empty] | T value |  | 3.7469 | 3.7469 |  | T value |  | 3.7469 | 3.7469
[9 empty] | MDL |  | 0.04961 | 1.2721 |  | MDL |  | 0.072074 | 4.5426
[9 empty] | LOQ |  | 0.1324 | 3.395 |  | LOQ |  | 0.19235 | 12.1235
[9 empty] | Error as % |  |  |  |  | Error as %
[9 empty] | Known/MDL |  |  |  |  | Known/MDL
[9 empty] | Mean PR |  |  |  |  | Mean PR
[9 empty] | S/N |  | 22.9153 | -0.8694 |  | S/N |  | 55.5227 | -1.9155 | [5 empty] | MEANS | -0.29516 | [9 empty] | MEANS | -2.3222
[5 empty] | Different from 0? |  | one sample T-test |  |  |  | p value | 0.4149 |  |  |  | p value | 0.3504
[9 empty] | count |  | 5 |  |  |  |  | 5
[9 empty] | std error of mean |  | 0.15183 |  |  |  |  | 0.54218
[9 empty] | df |  | 4 |  |  |  |  | 4
[9 empty] | t |  | 1.944 |  |  |  |  | 4.2831
[9 empty] | p value |  | 0.1238 |  |  |  |  | 0.012816
DIGESTION CHECK - is organic material completely converted to inorganic?
2026-02-27 | 12:48 | OM_2-27-2026_12-05-56PM.OMN | digest check | 24 | Unknown | 1 | 1 | TP | ug P/L | 0.0782 | 1.64 | 35.2 | TN | ug N/L | 0.549 | 7.13 | 370 |  | 1 | 1 |  | 35.2 | 35.2 | 35.2 | [6 empty] | 1 |  | 370 | 370 | 370 | [6 empty] | 42
2026-02-27 | 12:50 | OM_2-27-2026_12-05-56PM.OMN | digest check | 25 | Unknown | 1 | 1 | TP | ug P/L | 0.0765 | 1.63 | 35 | TN | ug N/L | 0.57 | 7.4 | 387 |  | 1 | 1 |  | 35 | 35 | 35 | [6 empty] | 1 |  | 387 | 387 | 387 | [6 empty] | 43
2026-02-27 | 12:51 | OM_2-27-2026_12-05-56PM.OMN | digest check | 26 | Unknown | 1 | 1 | TP | ug P/L | 0.0756 | 1.63 | 35 | TN | ug N/L | 0.56 | 7.21 | 376 |  | 1 | 1 |  | 35 | 35 | 35 | [6 empty] | 1 |  | 376 | 376 | 376 | [6 empty] | 44
2026-02-27 | 12:52 | OM_2-27-2026_12-05-56PM.OMN | digest check | 27 | Unknown | 1 | 1 | TP | ug P/L | 0.086 | 1.7 | 36.9 | TN | ug N/L | 0.541 | 6.98 | 361 |  | 1 | 1 |  | 36.9 | 36.9 | 36.9 | [6 empty] | 1 |  | 361 | 361 | 361 | [6 empty] | 45
2026-02-27 | 12:53 | OM_2-27-2026_12-05-56PM.OMN | digest check | 28 | Unknown | 1 | 1 | TP | ug P/L | 0.0793 | 1.69 | 36.8 | TN | ug N/L | 0.562 | 7.28 | 380 |  | 1 | 1 |  | 36.8 | 36.8 | 36.8 | [6 empty] | 1 |  | 380 | 380 | 380 | [6 empty] | 46
[11 empty] | Known Conc | 40 |  |  |  | Known Conc | 400 | [5 empty] | Known Conc | 40 | [9 empty] | Known Conc | 400
[11 empty] | Mean | 35.78 |  |  |  | Mean | 374.8 | [5 empty] | Mean | 35.78 | [9 empty] | Mean | 374.8
[11 empty] | Std Dev | 0.98082 |  |  |  | Std Dev | 9.8843 | [5 empty] | Std Dev | 0.98082 | [9 empty] | Std Dev | 9.8843
[11 empty] | CV or %RSD | 2.7412 |  |  |  | CV or %RSD | 2.6372 | [5 empty] | CV or %RSD | 2.7412 | [9 empty] | CV or %RSD | 2.6372
[11 empty] | T value | 3.7469 |  |  |  | T value | 3.7469 | [5 empty] | T value | 3.7469 | [9 empty] | T value | 3.7469
[11 empty] | MDL | 3.6751 |  |  |  | MDL | 37.0361 | [5 empty] | MDL | 3.6751 | [9 empty] | MDL | 37.0361
[11 empty] | LOQ | 9.8082 |  |  |  | LOQ | 98.8433 | [5 empty] | LOQ | 9.8082 | [9 empty] | LOQ | 98.8433
[11 empty] | Error as % | -10.55 |  |  |  | Error as % | -6.3 | [5 empty] | Error as % | -10.55 | [9 empty] | Error as % | -6.3
[11 empty] | Known/MDL |  |  |  |  | Known/MDL | [6 empty] | Known/MDL | [10 empty] | Known/MDL
[11 empty] | Mean PR | 89.45 |  |  |  | Mean PR | 93.7 | [5 empty] | Mean PR | [10 empty] | Mean PR
[11 empty] | S/N | 36.4798 |  |  |  | S/N | 37.9186 | [5 empty] | S/N | 36.4798 | [9 empty] | S/N | 37.9186
SPIKED BLANKS- used for MDL determination (250 ul spiking soln 1000/10000 into 10 ml sample)
2026-02-27 | 12:55 | OM_2-27-2026_12-05-56PM.OMN | spk blank | 29 | Unknown | 1 | 1 | TP | ug P/L | 0.051 | 1.09 | 20.2 | TN | ug N/L | 0.377 | 4.96 | 236 |  | 1 | 1 |  | 20.2 | 20.2 | 20.2 | [6 empty] | 1 |  | 236 | 236 | 236 | [6 empty] | 47
2026-02-27 | 12:56 | OM_2-27-2026_12-05-56PM.OMN | spk blank | 30 | Unknown | 1 | 1 | TP | ug P/L | 0.0516 | 1.12 | 21.1 | TN | ug N/L | 0.445 | 5.79 | 287 |  | 1 | 1 |  | 21.1 | 21.1 | 21.1 | [6 empty] | 1 |  | 287 | 287 | 287 | [6 empty] | 48
2026-02-27 | 12:57 | OM_2-27-2026_12-05-56PM.OMN | spk blank | 31 | Unknown | 1 | 1 | TP | ug P/L | 0.0519 | 1.15 | 21.8 | TN | ug N/L | 0.379 | 4.98 | 237 |  | 1 | 1 |  | 21.8 | 21.8 | 21.8 | [6 empty] | 1 |  | 237 | 237 | 237 | [6 empty] | 49
2026-02-27 | 12:59 | OM_2-27-2026_12-05-56PM.OMN | spk blank | 32 | Unknown | 1 | 1 | TP | ug P/L | 0.0488 | 1.06 | 19.6 | TN | ug N/L | 0.372 | 4.89 | 232 |  | 1 | 1 |  | 19.6 | 19.6 | 19.6 | [6 empty] | 1 |  | 232 | 232 | 232 | [6 empty] | 50
2026-02-27 | 13:00 | OM_2-27-2026_12-05-56PM.OMN | spk blank | 33 | Unknown | 1 | 1 | TP | ug P/L | 0.0488 | 1.07 | 19.7 | TN | ug N/L | 0.369 | 4.81 | 227 |  | 1 | 1 |  | 19.7 | 19.7 | 19.7 | [6 empty] | 1 |  | 227 | 227 | 227 | [6 empty] | 51
[11 empty] | Known Conc | 24.39 |  |  |  | Known Conc | 243.9 | [5 empty] | Known Conc | 24.39 | [9 empty] | Known Conc | 243.9
[11 empty] | Mean | 20.48 |  |  |  | Mean | 243.8 | [5 empty] | Mean | 20.48 | [9 empty] | Mean | 243.8
[11 empty] | Std Dev | 0.9471 |  |  |  | Std Dev | 24.4683 | [5 empty] | Std Dev | 0.9471 | [9 empty] | Std Dev | 24.4683
[11 empty] | CV or %RSD | 4.6245 |  |  |  | CV or %RSD | 10.0362 | [5 empty] | CV or %RSD | 4.6245 | [9 empty] | CV or %RSD | 10.0362
[11 empty] | T value | 3.7469 |  |  |  | T value | 3.7469 | [5 empty] | T value | 3.7469 | [9 empty] | T value | 3.7469
[11 empty] | MDL | 3.5487 |  |  |  | MDL | 91.6816 | [5 empty] | MDL | 3.5487 | [9 empty] | MDL | 91.6816
[11 empty] | LOQ | 9.471 |  |  |  | LOQ | 244.6835 | [5 empty] | LOQ | 9.471 | [9 empty] | LOQ | 244.6835
[11 empty] | Error as % | -16.0312 |  |  |  | Error as % | -0.041 | [5 empty] | Error as % | -16.0312 | [9 empty] | Error as % | -0.041
[11 empty] | Known/MDL |  |  |  |  | Known/MDL | [6 empty] | Known/MDL | [10 empty] | Known/MDL
[11 empty] | Mean PR | 83.9688 |  |  |  | Mean PR | 99.959 | [5 empty] | Mean PR | [10 empty] | Mean PR
[11 empty] | S/N | 21.6239 |  |  |  | S/N | 9.9639 | [5 empty] | S/N | 21.6239 | [9 empty] | S/N | 9.9639

## Raw Data
Detection Date | Detection Time | Run File Name | Sample ID | Sample Name | Cup Number | Sample Type | Auto Dilution Factor | Manual Dilution Factor | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Analyte Name | Concentration Units | Peak Height | Peak Area | Peak Concentration | Sample name
2026-02-27 | 12:07 | OM_2-27-2026_12-05-56PM.OMN | dye |  | S8 | Unknown | 1 | 1 | TP | ug P/L | 0.133 | 1.66 | 30.2 | TN | ug N/L | -0.00663 | -0.0863 | -113
2026-02-27 | 12:09 | OM_2-27-2026_12-05-56PM.OMN | dye |  | S8 | Unknown | 1 | 1 | TP | ug P/L | 0.133 | 1.66 | 30.4 | TN | ug N/L | 0.00637 | 0.155 | -87.1
2026-02-27 | 12:10 | OM_2-27-2026_12-05-56PM.OMN | NO2 |  | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.135 | 1.67 | 30.6 | TN | ug N/L | 1.98 | 26 | 2580
2026-02-27 | 12:11 | OM_2-27-2026_12-05-56PM.OMN | NO3 |  | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.134 | 1.58 | 27.7 | TN | ug N/L | 1.95 | 24.8 | 2460
2026-02-27 | 12:13 | OM_2-27-2026_12-05-56PM.OMN | NO2 |  | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.135 | 1.66 | 30.4 | TN | ug N/L | 1.88 | 24.6 | 2450
2026-02-27 | 12:14 | OM_2-27-2026_12-05-56PM.OMN | NO3 |  | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.139 | 1.73 | 32.4 | TN | ug N/L | 1.95 | 25 | 2480
2026-02-27 | 12:15 | OM_2-27-2026_12-05-56PM.OMN | water |  | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.138 | 1.69 | 31.2 | TN | ug N/L | 0.00705 | 0.16 | -86.6
2026-02-27 | 12:16 | OM_2-27-2026_12-05-56PM.OMN | water |  | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.134 | 1.63 | 29.5 | TN | ug N/L | 0.0153 | -0.0446 | -108
2026-02-27 | 12:18 | OM_2-27-2026_12-05-56PM.OMN | bisulfate carrier |  | S4 | Unknown | 1 | 1 | TP | ug P/L | 0.00979 | 0.114 | -16.4 | TN | ug N/L | -0.0156 | -0.17 | -121
2026-02-27 | 12:19 | OM_2-27-2026_12-05-56PM.OMN | bisulfate carrier |  | S4 | Unknown | 1 | 1 | TP | ug P/L | 0.0102 | 0.113 | -16.4 | TN | ug N/L | 0.0184 | 0.378 | -63.6
2026-02-27 | 12:20 | OM_2-27-2026_12-05-56PM.OMN | cal std 150 |  | 1 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.288 | 5.46 | 150 | TN | ug N/L | 1.89 | 24.1 | 1500
2026-02-27 | 12:22 | OM_2-27-2026_12-05-56PM.OMN | cal std 150 |  | 1 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.289 | 5.5 | 150 | TN | ug N/L | 1.95 | 24.8 | 1500
2026-02-27 | 12:23 | OM_2-27-2026_12-05-56PM.OMN | cal std 100 |  | 3 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.191 | 3.67 | 100 | TN | ug N/L | 1.36 | 17.4 | 1000
2026-02-27 | 12:24 | OM_2-27-2026_12-05-56PM.OMN | cal std 100 |  | 3 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.192 | 3.73 | 100 | TN | ug N/L | 1.31 | 16.8 | 1000
2026-02-27 | 12:25 | OM_2-27-2026_12-05-56PM.OMN | cal std 50 |  | 5 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0998 | 2.06 | 50 | TN | ug N/L | 0.717 | 9.25 | 500
2026-02-27 | 12:27 | OM_2-27-2026_12-05-56PM.OMN | cal std 50 |  | 5 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0992 | 2.24 | 50 | TN | ug N/L | 0.696 | 8.98 | 500
2026-02-27 | 12:28 | OM_2-27-2026_12-05-56PM.OMN | cal std 25 |  | 7 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0505 | 1.16 | 25 | TN | ug N/L | 0.402 | 5.26 | 250
2026-02-27 | 12:29 | OM_2-27-2026_12-05-56PM.OMN | cal std 25 |  | 7 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0501 | 1.13 | 25 | TN | ug N/L | 0.387 | 5.08 | 250
2026-02-27 | 12:30 | OM_2-27-2026_12-05-56PM.OMN | cal std 10 |  | 9 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0231 | 0.611 | 10 | TN | ug N/L | 0.19 | 2.61 | 100
2026-02-27 | 12:32 | OM_2-27-2026_12-05-56PM.OMN | cal std 10 |  | 9 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0333 | 0.828 | 10 | TN | ug N/L | 0.196 | 2.69 | 100
2026-02-27 | 12:33 | OM_2-27-2026_12-05-56PM.OMN | cal std 5 |  | 11 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0267 | 0.64 | 5 | TN | ug N/L | 0.154 | 2.17 | 50
2026-02-27 | 12:34 | OM_2-27-2026_12-05-56PM.OMN | cal std 5 |  | 11 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0188 | 0.495 | 5 | TN | ug N/L | 0.131 | 1.87 | 50
2026-02-27 | 12:36 | OM_2-27-2026_12-05-56PM.OMN | cal std 0 |  | 13 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.0113 | 0.274 | 0 | TN | ug N/L | 0.0691 | 1.07 | 0
2026-02-27 | 12:37 | OM_2-27-2026_12-05-56PM.OMN | cal std 0 |  | 13 | Calibration Standard | 1 | 1 | TP | ug P/L | 0.013 | 0.28 | 0 | TN | ug N/L | 0.0703 | 1.08 | 0
2026-02-27 | 12:38 | OM_2-27-2026_12-05-56PM.OMN | NO2 |  | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.137 | 1.72 | 37.5 | TN | ug N/L | 1.87 | 24.5 | 1500
2026-02-27 | 12:39 | OM_2-27-2026_12-05-56PM.OMN | NO3 |  | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.138 | 1.7 | 37 | TN | ug N/L | 1.98 | 25.1 | 1540
2026-02-27 | 12:41 | OM_2-27-2026_12-05-56PM.OMN | water |  | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.136 | 1.65 | 35.6 | TN | ug N/L | 0.00786 | 0.124 | -59.6
2026-02-27 | 12:42 | OM_2-27-2026_12-05-56PM.OMN | water blank |  | 19 | Unknown | 1 | 1 | TP | ug P/L | 0.0128 | 0.314 | -0.0151 | TN | ug N/L | 0.0703 | 1.1 | -0.261
2026-02-27 | 12:43 | OM_2-27-2026_12-05-56PM.OMN | water blank |  | 20 | Unknown | 1 | 1 | TP | ug P/L | 0.0111 | 0.291 | -0.624 | TN | ug N/L | 0.0676 | 1.05 | -3.38
2026-02-27 | 12:45 | OM_2-27-2026_12-05-56PM.OMN | water blank |  | 21 | Unknown | 1 | 1 | TP | ug P/L | 0.0122 | 0.313 | -0.0603 | TN | ug N/L | 0.0688 | 1.06 | -2.77
2026-02-27 | 12:46 | OM_2-27-2026_12-05-56PM.OMN | water blank |  | 22 | Unknown | 1 | 1 | TP | ug P/L | 0.0126 | 0.287 | -0.706 | TN | ug N/L | 0.069 | 1.06 | -2.88
2026-02-27 | 12:47 | OM_2-27-2026_12-05-56PM.OMN | water blank |  | 23 | Unknown | 1 | 1 | TP | ug P/L | 0.0124 | 0.312 | -0.0704 | TN | ug N/L | 0.0692 | 1.07 | -2.32
2026-02-27 | 12:48 | OM_2-27-2026_12-05-56PM.OMN | digest check |  | 24 | Unknown | 1 | 1 | TP | ug P/L | 0.0782 | 1.64 | 35.2 | TN | ug N/L | 0.549 | 7.13 | 370
2026-02-27 | 12:50 | OM_2-27-2026_12-05-56PM.OMN | digest check |  | 25 | Unknown | 1 | 1 | TP | ug P/L | 0.0765 | 1.63 | 35 | TN | ug N/L | 0.57 | 7.4 | 387
2026-02-27 | 12:51 | OM_2-27-2026_12-05-56PM.OMN | digest check |  | 26 | Unknown | 1 | 1 | TP | ug P/L | 0.0756 | 1.63 | 35 | TN | ug N/L | 0.56 | 7.21 | 376
2026-02-27 | 12:52 | OM_2-27-2026_12-05-56PM.OMN | digest check |  | 27 | Unknown | 1 | 1 | TP | ug P/L | 0.086 | 1.7 | 36.9 | TN | ug N/L | 0.541 | 6.98 | 361
2026-02-27 | 12:53 | OM_2-27-2026_12-05-56PM.OMN | digest check |  | 28 | Unknown | 1 | 1 | TP | ug P/L | 0.0793 | 1.69 | 36.8 | TN | ug N/L | 0.562 | 7.28 | 380
2026-02-27 | 12:55 | OM_2-27-2026_12-05-56PM.OMN | spk blank |  | 29 | Unknown | 1 | 1 | TP | ug P/L | 0.051 | 1.09 | 20.2 | TN | ug N/L | 0.377 | 4.96 | 236
2026-02-27 | 12:56 | OM_2-27-2026_12-05-56PM.OMN | spk blank |  | 30 | Unknown | 1 | 1 | TP | ug P/L | 0.0516 | 1.12 | 21.1 | TN | ug N/L | 0.445 | 5.79 | 287
2026-02-27 | 12:57 | OM_2-27-2026_12-05-56PM.OMN | spk blank |  | 31 | Unknown | 1 | 1 | TP | ug P/L | 0.0519 | 1.15 | 21.8 | TN | ug N/L | 0.379 | 4.98 | 237
2026-02-27 | 12:59 | OM_2-27-2026_12-05-56PM.OMN | spk blank |  | 32 | Unknown | 1 | 1 | TP | ug P/L | 0.0488 | 1.06 | 19.6 | TN | ug N/L | 0.372 | 4.89 | 232
2026-02-27 | 13:00 | OM_2-27-2026_12-05-56PM.OMN | spk blank |  | 33 | Unknown | 1 | 1 | TP | ug P/L | 0.0488 | 1.07 | 19.7 | TN | ug N/L | 0.369 | 4.81 | 227
2026-02-27 | 13:01 | OM_2-27-2026_12-05-56PM.OMN | S1 | B4 6/18/25 | 34 | Unknown | 1 | 1 | TP | ug P/L | 0.523 | 9.65 | 307 | TN | ug N/L | 1.46 | 18.5 | 1100
2026-02-27 | 13:02 | OM_2-27-2026_12-05-56PM.OMN | S2 | C2 7/03/25 | 35 | Unknown | 1 | 1 | TP | ug P/L | 0.146 | 2.84 | 69.7 | TN | ug N/L | 0.885 | 11.3 | 634
2026-02-27 | 13:04 | OM_2-27-2026_12-05-56PM.OMN | S3 | C1 6/18/25 | 36 | Unknown | 1 | 1 | TP | ug P/L | 0.854 | 15.5 | 568 | TN | ug N/L | 2.37 | 30.1 | 1880
2026-02-27 | 13:05 | OM_2-27-2026_12-05-56PM.OMN | S4 | D4 8/21/25 | 37 | Unknown | 1 | 1 | TP | ug P/L | 0.0558 | 1.2 | 23.2 | TN | ug N/L | 1.13 | 14.4 | 834
2026-02-27 | 13:06 | OM_2-27-2026_12-05-56PM.OMN | S5 | 704 9/15/25 | 38 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.119 | 2.35 | 55.3 | TN | ug N/L | 1.16 | 14.8 | 861
2026-02-27 | 13:07 | OM_2-27-2026_12-05-56PM.OMN | DUP S5 | 704 DUP 9/15/25 | 44 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.112 | 2.44 | 57.9 | TN | ug N/L | 1.12 | 14.3 | 825
2026-02-27 | 13:09 | OM_2-27-2026_12-05-56PM.OMN | S6 | D3 7/31/25 | 39 | Unknown | 1 | 1 | TP | ug P/L | 0.226 | 4.31 | 115 | TN | ug N/L | 1.34 | 17.1 | 1000
2026-02-27 | 13:10 | OM_2-27-2026_12-05-56PM.OMN | S7 | C1 7/31/25 | 40 | Unknown | 1 | 1 | TP | ug P/L | 0.14 | 2.74 | 66.8 | TN | ug N/L | 1 | 12.9 | 733
2026-02-27 | 13:11 | OM_2-27-2026_12-05-56PM.OMN | S8 | B4 7/31/25 | 41 | Unknown | 1 | 1 | TP | ug P/L | 0.0824 | 1.79 | 39.5 | TN | ug N/L | 0.662 | 8.44 | 453
2026-02-27 | 13:13 | OM_2-27-2026_12-05-56PM.OMN | S9 | 706 8/07/25 | 42 | Unknown | 1 | 1 | TP | ug P/L | 0.0595 | 1.31 | 26.2 | TN | ug N/L | 0.547 | 7.04 | 365
2026-02-27 | 13:14 | OM_2-27-2026_12-05-56PM.OMN | S10 | R1 10/27/25 | 43 | Spike Set | 1 | 1 | TP | ug P/L | 0.955 | 16.3 | 606 | TN | ug N/L | 1.14 | 14.2 | 816
2026-02-27 | 13:15 | OM_2-27-2026_12-05-56PM.OMN | SPK S10 |  | 45 | Spike Set | 1 | 1 | TP | ug P/L | 0.923 | 16.7 | 627 | TN | ug N/L | 1.33 | 16.7 | 983
2026-02-27 | 13:16 | OM_2-27-2026_12-05-56PM.OMN | water |  | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.139 | 1.51 | 31.7 | TN | ug N/L | 0.00601 | 0.0787 | -62.4
2026-02-27 | 13:18 | OM_2-27-2026_12-05-56PM.OMN | CHECK25 |  | 15 | Check Standard | 1 | 1 | TP | ug P/L | 0.0521 | 1.17 | 22.6 | TN | ug N/L | 0.366 | 4.8 | 226
2026-02-27 | 13:19 | OM_2-27-2026_12-05-56PM.OMN | S11 | R2 6/18/25 | 46 | Unknown | 1 | 1 | TP | ug P/L | 0.162 | 3.08 | 76.7 | TN | ug N/L | 0.354 | 4.6 | 214
2026-02-27 | 13:20 | OM_2-27-2026_12-05-56PM.OMN | S12 | C2 6/18/25 | 47 | Unknown | 1 | 1 | TP | ug P/L | 1.04 | 18.7 | 736 | TN | ug N/L | 2.24 | 28.2 | 1750
2026-02-27 | 13:22 | OM_2-27-2026_12-05-56PM.OMN | S13 | B3 8/21/25 | 48 | Unknown | 1 | 1 | TP | ug P/L | 0.11 | 2.21 | 51.4 | TN | ug N/L | 1.35 | 17.1 | 1010
2026-02-27 | 13:23 | OM_2-27-2026_12-05-56PM.OMN | S14 | diospryus virginiana 8/18/25 | 49 | Unknown | 1 | 1 | TP | ug P/L | 0.0914 | 1.73 | 37.7 | TN | ug N/L | 0.33 | 4.33 | 197
2026-02-27 | 13:24 | OM_2-27-2026_12-05-56PM.OMN | S15 | D1 8/21/25 | 50 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.163 | 3.18 | 79.8 | TN | ug N/L | 1.23 | 15.6 | 910
2026-02-27 | 13:25 | OM_2-27-2026_12-05-56PM.OMN | DUP S15 |  | 56 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.171 | 3.36 | 85.3 | TN | ug N/L | 1.24 | 15.9 | 929
2026-02-27 | 13:27 | OM_2-27-2026_12-05-56PM.OMN | S16 | C3 8/21/25 | 51 | Unknown | 1 | 1 | TP | ug P/L | 0.0577 | 1.27 | 25.2 | TN | ug N/L | 0.757 | 9.69 | 532
2026-02-27 | 13:28 | OM_2-27-2026_12-05-56PM.OMN | S17 | 701 9/01/25 | 52 | Unknown | 1 | 1 | TP | ug P/L | 0.116 | 2.28 | 53.4 | TN | ug N/L | 1.28 | 16.3 | 953
2026-02-27 | 13:29 | OM_2-27-2026_12-05-56PM.OMN | S18 | B2 7/31/25 | 53 | Unknown | 1 | 1 | TP | ug P/L | 0.0692 | 1.43 | 29.5 | TN | ug N/L | 0.659 | 8.46 | 454
2026-02-27 | 13:30 | OM_2-27-2026_12-05-56PM.OMN | S19 | D3 9/01/25 | 54 | Unknown | 1 | 1 | TP | ug P/L | 0.226 | 4.29 | 114 | TN | ug N/L | 1.37 | 17.4 | 1020
2026-02-27 | 13:32 | OM_2-27-2026_12-05-56PM.OMN | S20 | A2 8/21/25 | 55 | Spike Set | 1 | 1 | TP | ug P/L | 0.102 | 2.08 | 47.6 | TN | ug N/L | 0.67 | 8.75 | 472
2026-02-27 | 13:33 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S20 |  | 57 | Spike Set | 1 | 1 | TP | ug P/L | 0.149 | 2.89 | 71.1 | TN | ug N/L | 0.947 | 12.1 | 685
2026-02-27 | 13:34 | OM_2-27-2026_12-05-56PM.OMN | bisulfate |  | S4 | Unknown | 1 | 1 | TP | ug P/L | 0.00835 | 0.0885 | -5.76 | TN | ug N/L | 0.0202 | 0.397 | -43.1
2026-02-27 | 13:36 | OM_2-27-2026_12-05-56PM.OMN | CHECK25 |  | 16 | Check Standard | 1 | 1 | TP | ug P/L | 0.0593 | 1.28 | 25.4 | TN | ug N/L | 0.435 | 5.62 | 277
2026-02-27 | 13:37 | OM_2-27-2026_12-05-56PM.OMN | NO2 |  | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.138 | 1.73 | 37.9 | TN | ug N/L | 1.94 | 25.1 | 1540
2026-02-27 | 13:38 | OM_2-27-2026_12-05-56PM.OMN | NO3 |  | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.135 | 1.68 | 36.5 | TN | ug N/L | 2.03 | 25.8 | 1580
2026-02-27 | 13:39 | OM_2-27-2026_12-05-56PM.OMN | S21 | R1 7/03/25 | 58 | Unknown | 1 | 1 | TP | ug P/L | 0.374 | 6.95 | 204 | TN | ug N/L | 0.582 | 7.64 | 402
2026-02-27 | 13:41 | OM_2-27-2026_12-05-56PM.OMN | S22 | B3 6/18/25 | 59 | Unknown | 1 | 1 | TP | ug P/L | 0.495 | 9.05 | 283 | TN | ug N/L | 1.51 | 15.5 | 902
2026-02-27 | 13:42 | OM_2-27-2026_12-05-56PM.OMN | S23 | B2 9/29/25 | 60 | Unknown | 1 | 1 | TP | ug P/L | 0.198 | 3.77 | 97.8 | TN | ug N/L | 0.946 | 12.1 | 686
2026-02-27 | 13:43 | OM_2-27-2026_12-05-56PM.OMN | S24 | C4 9/29/25 | 61 | Unknown | 1 | 1 | TP | ug P/L | 0.296 | 5.56 | 156 | TN | ug N/L | 3.25 | 41.4 | 2680
2026-02-27 | 13:45 | OM_2-27-2026_12-05-56PM.OMN | S25 | C2 9/29/25 | 62 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.178 | 3.38 | 86 | TN | ug N/L | 1.24 | 15.7 | 913
2026-02-27 | 13:46 | OM_2-27-2026_12-05-56PM.OMN | DUP S25 | C2 9/29/25 | 68 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.174 | 3.36 | 85.3 | TN | ug N/L | 1.23 | 15.5 | 905
2026-02-27 | 13:47 | OM_2-27-2026_12-05-56PM.OMN | S26 | D3 9/29/25 | 63 | Unknown | 1 | 1 | TP | ug P/L | 0.348 | 6.47 | 187 | TN | ug N/L | 2.09 | 26.5 | 1630
2026-02-27 | 13:48 | OM_2-27-2026_12-05-56PM.OMN | S27 | A4 7/03/25 | 64 | Unknown | 1 | 1 | TP | ug P/L | 0.106 | 2.14 | 49.5 | TN | ug N/L | 0.587 | 7.55 | 397
2026-02-27 | 13:50 | OM_2-27-2026_12-05-56PM.OMN | S28 | C4 8/21/25 | 65 | Unknown | 1 | 1 | TP | ug P/L | 0.0996 | 2.06 | 47.1 | TN | ug N/L | 1.16 | 14.8 | 858
2026-02-27 | 13:51 | OM_2-27-2026_12-05-56PM.OMN | S29 | 706 9/15/25 | 66 | Unknown | 1 | 1 | TP | ug P/L | 0.113 | 2.24 | 52.1 | TN | ug N/L | 1.11 | 14.1 | 815
2026-02-27 | 13:52 | OM_2-27-2026_12-05-56PM.OMN | S30 | C4 9/01/25 | 67 | Spike Set | 1 | 1 | TP | ug P/L | 0.0637 | 1.35 | 27.3 | TN | ug N/L | 0.976 | 12.5 | 709
2026-02-27 | 13:53 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S30 |  | 69 | Spike Set | 1 | 1 | TP | ug P/L | 0.13 | 2.34 | 55.1 | TN | ug N/L | 1.48 | 18.7 | 1110
2026-02-27 | 13:55 | OM_2-27-2026_12-05-56PM.OMN | water |  | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.157 | 1.89 | 42.2 | TN | ug N/L | 0.00555 | -0.115 | -74
2026-02-27 | 13:56 | OM_2-27-2026_12-05-56PM.OMN | CHECK25 |  | 17 | Check Standard | 1 | 1 | TP | ug P/L | 0.0509 | 1.16 | 22.1 | TN | ug N/L | 0.374 | 4.88 | 231
2026-02-27 | 13:57 | OM_2-27-2026_12-05-56PM.OMN | S31 | B2 7/03/25 | 70 | Unknown | 1 | 1 | TP | ug P/L | 0.0914 | 1.86 | 41.4 | TN | ug N/L | 0.809 | 10.4 | 575
2026-02-27 | 13:59 | OM_2-27-2026_12-05-56PM.OMN | S32 | R1 7/31/25 | 71 | Unknown | 1 | 1 | TP | ug P/L | 0.487 | 8.97 | 280 | TN | ug N/L | 0.756 | 9.56 | 523
2026-02-27 | 14:00 | OM_2-27-2026_12-05-56PM.OMN | S33 | B3 9/01/25 | 72 | Unknown | 1 | 1 | TP | ug P/L | 0.0946 | 1.95 | 43.9 | TN | ug N/L | 0.9 | 11.4 | 641
2026-02-27 | 14:01 | OM_2-27-2026_12-05-56PM.OMN | S34 | D4 9/01/25 | 73 | Unknown | 1 | 1 | TP | ug P/L | 0.102 | 2.04 | 46.5 | TN | ug N/L | 1.85 | 23.5 | 1430
2026-02-27 | 14:02 | OM_2-27-2026_12-05-56PM.OMN | S35 | C3 7/03/25 | 74 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.242 | 4.61 | 124 | TN | ug N/L | 1.12 | 14.2 | 821
2026-02-27 | 14:04 | OM_2-27-2026_12-05-56PM.OMN | DUP S35 | C3 7/03/25 | 80 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.251 | 4.74 | 129 | TN | ug N/L | 1.16 | 14.7 | 853
2026-02-27 | 14:05 | OM_2-27-2026_12-05-56PM.OMN | S36 | C3 10/27/25 | 75 | Unknown | 1 | 1 | TP | ug P/L | 0.0883 | 1.84 | 40.8 | TN | ug N/L | 0.923 | 11.8 | 665
2026-02-27 | 14:06 | OM_2-27-2026_12-05-56PM.OMN | S37 | A1 6/18/25 | 76 | Unknown | 1 | 1 | TP | ug P/L | 0.614 | 11.3 | 374 | TN | ug N/L | 1.51 | 19.2 | 1140
2026-02-27 | 14:08 | OM_2-27-2026_12-05-56PM.OMN | S38 | 702 8/7/25 | 77 | Unknown | 1 | 1 | TP | ug P/L | 0.104 | 2.08 | 47.7 | TN | ug N/L | 0.799 | 10.2 | 566
2026-02-27 | 14:09 | OM_2-27-2026_12-05-56PM.OMN | S39 | D3 7/03/25 | 78 | Unknown | 1 | 1 | TP | ug P/L | 0.859 | 15.8 | 583 | TN | ug N/L | 3.17 | 40 | 2580
2026-02-27 | 14:10 | OM_2-27-2026_12-05-56PM.OMN | S40 | C3 7/31/25 | 79 | Spike Set | 1 | 1 | TP | ug P/L | 0.0597 | 1.3 | 26.1 | TN | ug N/L | 0.723 | 9.28 | 505
2026-02-27 | 14:11 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S40 |  | 81 | Spike Set | 1 | 1 | TP | ug P/L | 0.106 | 2.14 | 49.3 | TN | ug N/L | 0.973 | 12.4 | 704
2026-02-27 | 14:13 | OM_2-27-2026_12-05-56PM.OMN | BISULFATE |  | S4 | Unknown | 1 | 1 | TP | ug P/L | 0.00831 | 0.083 | -5.9 | TN | ug N/L | 0.0192 | 0.383 | -43.9
2026-02-27 | 14:14 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 |  | 18 | Check Standard | 1 | 1 | TP | ug P/L | 0.0511 | 1.13 | 21.5 | TN | ug N/L | 0.376 | 4.97 | 236
2026-02-27 | 14:15 | OM_2-27-2026_12-05-56PM.OMN | NO2 |  | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.135 | 1.64 | 35.4 | TN | ug N/L | 1.83 | 24 | 1460
2026-02-27 | 14:17 | OM_2-27-2026_12-05-56PM.OMN | NO3 |  | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.139 | 1.72 | 37.5 | TN | ug N/L | 1.97 | 25 | 1530
2026-02-27 | 14:18 | OM_2-27-2026_12-05-56PM.OMN | S41 | B3 7/03/25 | 82 | Unknown | 1 | 1 | TP | ug P/L | 0.106 | 2.16 | 49.8 | TN | ug N/L | 0.823 | 10.6 | 591
2026-02-27 | 14:19 | OM_2-27-2026_12-05-56PM.OMN | S42 | R2 7/03/25 | 83 | Unknown | 1 | 1 | TP | ug P/L | 0.216 | 3.82 | 99.3 | TN | ug N/L | 0.346 | 3.98 | 176
2026-02-27 | 14:20 | OM_2-27-2026_12-05-56PM.OMN | S43 | D1 9/29/25 | 84 | Unknown | 1 | 1 | TP | ug P/L | 0.378 | 6.81 | 199 | TN | ug N/L | 2.86 | 36.3 | 2310
2026-02-27 | 14:22 | OM_2-27-2026_12-05-56PM.OMN | S44 | 703 8/07/25 | 85 | Unknown | 1 | 1 | TP | ug P/L | 0.0606 | 1.31 | 26.3 | TN | ug N/L | 0.541 | 7 | 363
2026-02-27 | 14:23 | OM_2-27-2026_12-05-56PM.OMN | S45 | D1 10/27/25 | 86 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.103 | 2.05 | 46.8 | TN | ug N/L | 0.956 | 12.2 | 692
2026-02-27 | 14:24 | OM_2-27-2026_12-05-56PM.OMN | DUP S45 | D1 10/27/25 | 92 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.101 | 2.01 | 45.8 | TN | ug N/L | 1.06 | 13.6 | 778
2026-02-27 | 14:25 | OM_2-27-2026_12-05-56PM.OMN | S46 | A2 9/29/25 | 87 | Unknown | 1 | 1 | TP | ug P/L | 0.287 | 5.33 | 148 | TN | ug N/L | 1.48 | 18.8 | 1120
2026-02-27 | 14:27 | OM_2-27-2026_12-05-56PM.OMN | S47 | alnus serrulata (shade) 8/18/25 | 88 | Unknown | 1 | 1 | TP | ug P/L | 0.119 | 2.39 | 56.6 | TN | ug N/L | 0.381 | 4.97 | 237
2026-02-27 | 14:28 | OM_2-27-2026_12-05-56PM.OMN | S48 | C4 7/03/25 | 89 | Unknown | 1 | 1 | TP | ug P/L | 0.111 | 2.22 | 51.6 | TN | ug N/L | 0.814 | 10.4 | 577
2026-02-27 | 14:29 | OM_2-27-2026_12-05-56PM.OMN | S49 | robina pseudoacacia 8/18/25 | 90 | Unknown | 1 | 1 | TP | ug P/L | 0.166 | 2.71 | 65.8 | TN | ug N/L | 1.58 | 20.1 | 1210
2026-02-27 | 14:31 | OM_2-27-2026_12-05-56PM.OMN | S50 | A3 7/31/25 | 91 | Spike Set | 1 | 1 | TP | ug P/L | 0.13 | 2.57 | 61.7 | TN | ug N/L | 1.05 | 13.4 | 769
2026-02-27 | 14:32 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S50 | A3 7/31/25 | 93 | Spike Set | 1 | 1 | TP | ug P/L | 0.178 | 3.35 | 85 | TN | ug N/L | 1.32 | 16.8 | 985
2026-02-27 | 14:33 | OM_2-27-2026_12-05-56PM.OMN | BLANK |  | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.135 | 1.63 | 35 | TN | ug N/L | 0.00585 | 0.114 | -60.2
2026-02-27 | 14:34 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 |  | 15 | Check Standard | 1 | 1 | TP | ug P/L | 0.0518 | 1.12 | 21.2 | TN | ug N/L | 0.398 | 4.92 | 233
2026-02-27 | 14:36 | OM_2-27-2026_12-05-56PM.OMN | S51 | A4 8/21/25 | 94 | Unknown | 1 | 1 | TP | ug P/L | 0.0488 | 1.14 | 21.6 | TN | ug N/L | 0.564 | 7.17 | 373
2026-02-27 | 14:37 | OM_2-27-2026_12-05-56PM.OMN | S52 | A2 10/27/25 | 95 | Unknown | 1 | 1 | TP | ug P/L | 0.157 | 3.03 | 75.3 | TN | ug N/L | 0.957 | 12.2 | 693
2026-02-27 | 14:38 | OM_2-27-2026_12-05-56PM.OMN | S53 | C4 10/27/25 | 96 | Unknown | 1 | 1 | TP | ug P/L | 0.282 | 5.28 | 146 | TN | ug N/L | 1.58 | 20 | 1200
2026-02-27 | 14:40 | OM_2-27-2026_12-05-56PM.OMN | S54 | R2 7/31/25 | 97 | Unknown | 1 | 1 | TP | ug P/L | 0.23 | 4.4 | 117 | TN | ug N/L | 0.35 | 4.53 | 210
2026-02-27 | 14:41 | OM_2-27-2026_12-05-56PM.OMN | S55 | 703 7/10/25 | 98 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.0587 | 1.26 | 24.9 | TN | ug N/L | 0.629 | 8.15 | 434
2026-02-27 | 14:42 | OM_2-27-2026_12-05-56PM.OMN | DUP S55 | 703 7/10/25 | 104 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.0553 | 1.2 | 23.2 | TN | ug N/L | 0.678 | 8.31 | 444
2026-02-27 | 14:43 | OM_2-27-2026_12-05-56PM.OMN | S56 | C3 9/01/25 | 99 | Unknown | 1 | 1 | TP | ug P/L | 0.0578 | 1.27 | 25.2 | TN | ug N/L | 0.963 | 12.2 | 692
2026-02-27 | 14:45 | OM_2-27-2026_12-05-56PM.OMN | S57 | B2 8/21/25 | 100 | Unknown | 1 | 1 | TP | ug P/L | 0.143 | 2.73 | 66.5 | TN | ug N/L | 1.16 | 14.7 | 851
2026-02-27 | 14:46 | OM_2-27-2026_12-05-56PM.OMN | S58 | A1 8/21/25 | 101 | Unknown | 1 | 1 | TP | ug P/L | 0.111 | 2.27 | 53 | TN | ug N/L | 0.682 | 8.77 | 474
2026-02-27 | 14:47 | OM_2-27-2026_12-05-56PM.OMN | S59 | 701 7/10/25 | 102 | Unknown | 1 | 1 | TP | ug P/L | 0.0892 | 1.82 | 40.4 | TN | ug N/L | 0.94 | 12 | 679
2026-02-27 | 14:48 | OM_2-27-2026_12-05-56PM.OMN | S60 | A4 7/31/25 | 103 | Spike Set | 1 | 1 | TP | ug P/L | 0.0639 | 1.37 | 28 | TN | ug N/L | 0.604 | 7.78 | 411
2026-02-27 | 14:50 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S60 |  | 105 | Spike Set | 1 | 1 | TP | ug P/L | 0.11 | 2.26 | 52.7 | TN | ug N/L | 0.892 | 11.4 | 638
2026-02-27 | 14:51 | OM_2-27-2026_12-05-56PM.OMN | BISULFATE |  | S4 | Unknown | 1 | 1 | TP | ug P/L | 0.0178 | 0.216 | -2.53 | TN | ug N/L | 0.0165 | 0.244 | -52.4
2026-02-27 | 14:52 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 |  | 16 | Check Standard | 1 | 1 | TP | ug P/L | 0.0503 | 1.14 | 21.6 | TN | ug N/L | 0.371 | 4.84 | 229
2026-02-27 | 14:54 | OM_2-27-2026_12-05-56PM.OMN | NO2 |  | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.134 | 1.63 | 35.1 | TN | ug N/L | 1.87 | 24.4 | 1490
2026-02-27 | 14:55 | OM_2-27-2026_12-05-56PM.OMN | NO3 |  | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.135 | 1.61 | 34.6 | TN | ug N/L | 1.94 | 24.6 | 1510
2026-02-27 | 14:56 | OM_2-27-2026_12-05-56PM.OMN | S61 | A3 7/03/25 | 106 | Unknown | 1 | 1 | TP | ug P/L | 0.145 | 2.84 | 69.6 | TN | ug N/L | 1.03 | 13.1 | 748
2026-02-27 | 14:57 | OM_2-27-2026_12-05-56PM.OMN | S62 | R2 8/18/25 | 107 | Unknown | 1 | 1 | TP | ug P/L | 0.0886 | 1.8 | 39.7 | TN | ug N/L | 0.373 | 4.88 | 231
2026-02-27 | 14:59 | OM_2-27-2026_12-05-56PM.OMN | S63 | A4 9/01/25 | 108 | Unknown | 1 | 1 | TP | ug P/L | 0.101 | 1.91 | 42.8 | TN | ug N/L | 1.01 | 12.8 | 732
2026-02-27 | 15:00 | OM_2-27-2026_12-05-56PM.OMN | S64 | 705 9/15/25 | 109 | Unknown | 1 | 1 | TP | ug P/L | 0.0809 | 1.71 | 37.3 | TN | ug N/L | 0.967 | 12.3 | 698
2026-02-27 | 15:01 | OM_2-27-2026_12-05-56PM.OMN | S65 | C1 10/27/25 | 110 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.107 | 2.1 | 48.3 | TN | ug N/L | 1.23 | 15.7 | 916
2026-02-27 | 15:03 | OM_2-27-2026_12-05-56PM.OMN | DUP S65 |  | 116 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.0996 | 2.03 | 46.3 | TN | ug N/L | 1.04 | 13.2 | 756
2026-02-27 | 15:04 | OM_2-27-2026_12-05-56PM.OMN | S66 | 707 7/10/25 | 111 | Unknown | 1 | 1 | TP | ug P/L | 0.114 | 2.31 | 54.2 | TN | ug N/L | 0.797 | 10.2 | 564
2026-02-27 | 15:05 | OM_2-27-2026_12-05-56PM.OMN | S67 | A3 9/01/25 | 112 | Unknown | 1 | 1 | TP | ug P/L | 0.153 | 2.97 | 73.6 | TN | ug N/L | 1.37 | 17.4 | 1020
2026-02-27 | 15:06 | OM_2-27-2026_12-05-56PM.OMN | S68 | B3 7/31`/25 | 113 | Unknown | 1 | 1 | TP | ug P/L | 0.0676 | 1.42 | 29.4 | TN | ug N/L | 0.68 | 8.69 | 469
2026-02-27 | 15:08 | OM_2-27-2026_12-05-56PM.OMN | S69 | B1 8/21/25 | 114 | Unknown | 1 | 1 | TP | ug P/L | 0.172 | 3.36 | 85.1 | TN | ug N/L | 1.11 | 14 | 808
2026-02-27 | 15:09 | OM_2-27-2026_12-05-56PM.OMN | S70 | B1 7/31/25 | 115 | Spike Set | 1 | 1 | TP | ug P/L | 0.0866 | 1.82 | 40.3 | TN | ug N/L | 0.768 | 9.87 | 543
2026-02-27 | 15:10 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S70 |  | 117 | Spike Set | 1 | 1 | TP | ug P/L | 0.14 | 2.76 | 67.2 | TN | ug N/L | 1.13 | 14.5 | 836
2026-02-27 | 15:12 | OM_2-27-2026_12-05-56PM.OMN | CHECK 150 |  | 2 | Check Standard | 1 | 1 | TP | ug P/L | 0.291 | 5.47 | 153 | TN | ug N/L | 1.91 | 24.2 | 1480
2026-02-27 | 15:13 | OM_2-27-2026_12-05-56PM.OMN | CHECK 100 |  | 4 | Check Standard | 1 | 1 | TP | ug P/L | 0.223 | 3.76 | 97.4 | TN | ug N/L | 1.35 | 17.1 | 1010
2026-02-27 | 15:14 | OM_2-27-2026_12-05-56PM.OMN | CHECK 50 |  | 6 | Check Standard | 1 | 1 | TP | ug P/L | 0.0967 | 1.97 | 44.6 | TN | ug N/L | 0.709 | 9.11 | 495
2026-02-27 | 15:15 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 |  | 8 | Check Standard | 1 | 1 | TP | ug P/L | 0.0515 | 1.13 | 21.5 | TN | ug N/L | 0.383 | 4.99 | 238
2026-02-27 | 15:17 | OM_2-27-2026_12-05-56PM.OMN | CHECK 10 |  | 10 | Check Standard | 1 | 1 | TP | ug P/L | 0.0255 | 0.717 | 10.4 | TN | ug N/L | 0.192 | 2.58 | 89.8
2026-02-27 | 15:18 | OM_2-27-2026_12-05-56PM.OMN | CHECK 5 |  | 12 | Check Standard | 1 | 1 | TP | ug P/L | 0.0147 | 0.42 | 2.7 | TN | ug N/L | 0.132 | 1.86 | 45.4
2026-02-27 | 15:19 | OM_2-27-2026_12-05-56PM.OMN | CHECK 0 |  | 14 | Check Standard | 1 | 1 | TP | ug P/L | 0.0115 | 0.25 | -1.66 | TN | ug N/L | 0.0703 | 1.06 | -2.71
2026-02-27 | 15:20 | OM_2-27-2026_12-05-56PM.OMN | S71 | D4 9/29/25 | 118 | Unknown | 1 | 1 | TP | ug P/L | 0.243 | 4.65 | 125 | TN | ug N/L | 3.12 | 39.9 | 2570
2026-02-27 | 15:22 | OM_2-27-2026_12-05-56PM.OMN | S72 | A2 7/03/25 | 119 | Unknown | 1 | 1 | TP | ug P/L | 0.118 | 2.42 | 57.4 | TN | ug N/L | 0.474 | 6.15 | 310
2026-02-27 | 15:23 | OM_2-27-2026_12-05-56PM.OMN | S73 | R2 8/21/25 | 120 | Unknown | 1 | 1 | TP | ug P/L | 0.204 | 3.91 | 102 | TN | ug N/L | 0.301 | 3.91 | 171
2026-02-27 | 15:24 | OM_2-27-2026_12-05-56PM.OMN | S74 | D4 7/03/25 | 121 | Unknown | 1 | 1 | TP | ug P/L | 0.828 | 15.3 | 556 | TN | ug N/L | 3.24 | 41.2 | 2660
2026-02-27 | 15:26 | OM_2-27-2026_12-05-56PM.OMN | S75 | R1 8/21/25 | 122 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.286 | 5.5 | 154 | TN | ug N/L | 0.432 | 5.56 | 273
2026-02-27 | 15:27 | OM_2-27-2026_12-05-56PM.OMN | DUP S75 |  | 128 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.26 | 4.95 | 136 | TN | ug N/L | 0.4 | 5.16 | 248
2026-02-27 | 15:28 | OM_2-27-2026_12-05-56PM.OMN | S76 | B2 9/01/25 | 123 | Unknown | 1 | 1 | TP | ug P/L | 0.186 | 3.67 | 94.6 | TN | ug N/L | 1.49 | 18.9 | 1130
2026-02-27 | 15:29 | OM_2-27-2026_12-05-56PM.OMN | S77 | R2 9/01/25 | 124 | Unknown | 1 | 1 | TP | ug P/L | 0.125 | 2.52 | 60.1 | TN | ug N/L | 0.227 | 3.02 | 116
2026-02-27 | 15:31 | OM_2-27-2026_12-05-56PM.OMN | S78 | 705 7/10.25 | 125 | Unknown | 1 | 1 | TP | ug P/L | 0.097 | 1.62 | 34.8 | TN | ug N/L | 0.928 | 11.8 | 667
2026-02-27 | 15:32 | OM_2-27-2026_12-05-56PM.OMN | S79 | 707 8/07/25 | 126 | Unknown | 1 | 1 | TP | ug P/L | 0.402 | 7.67 | 230 | TN | ug N/L | 1.37 | 17.5 | 1030
2026-02-27 | 15:33 | OM_2-27-2026_12-05-56PM.OMN | S80 | alnus serrulata (sun) 8/18/25 | 127 | Spike Set | 1 | 1 | TP | ug P/L | 0.314 | 5.94 | 169 | TN | ug N/L | 0.281 | 3.8 | 164
2026-02-27 | 15:35 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S80 |  | 129 | Spike Set | 1 | 1 | TP | ug P/L | 0.433 | 6.88 | 202 | TN | ug N/L | 0.555 | 6.65 | 340
2026-02-27 | 15:36 | OM_2-27-2026_12-05-56PM.OMN | BISULFATE |  | S4 | Unknown | 1 | 1 | TP | ug P/L | 0.00889 | 0.0925 | -5.66 | TN | ug N/L | 0.0187 | 0.376 | -44.4
2026-02-27 | 15:37 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 |  | 17 | Check Standard | 1 | 1 | TP | ug P/L | 0.0534 | 1.12 | 21.1 | TN | ug N/L | 0.376 | 4.9 | 232
2026-02-27 | 15:38 | OM_2-27-2026_12-05-56PM.OMN | NO2 |  | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.158 | 2.12 | 48.8 | TN | ug N/L | 1.86 | 24.3 | 1480
2026-02-27 | 15:40 | OM_2-27-2026_12-05-56PM.OMN | NO3 |  | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.142 | 1.97 | 44.5 | TN | ug N/L | 1.64 | 20.9 | 1260
2026-02-27 | 15:41 | OM_2-27-2026_12-05-56PM.OMN | S81 | 704 8/07/25 | 130 | Unknown | 1 | 1 | TP | ug P/L | 0.0783 | 1.74 | 38.1 | TN | ug N/L | 0.824 | 10.4 | 576
2026-02-27 | 15:42 | OM_2-27-2026_12-05-56PM.OMN | S82 | 703 9/15/25 | 131 | Unknown | 1 | 1 | TP | ug P/L | 0.0991 | 1.99 | 45.1 | TN | ug N/L | 1.15 | 14.7 | 851
2026-02-27 | 15:44 | OM_2-27-2026_12-05-56PM.OMN | S83 | C1 7/03/25 | 132 | Unknown | 1 | 1 | TP | ug P/L | 0.19 | 3.6 | 92.5 | TN | ug N/L | 1.89 | 23.9 | 1460
2026-02-27 | 15:45 | OM_2-27-2026_12-05-56PM.OMN | S84 | 706 7/10/25 | 133 | Unknown | 1 | 1 | TP | ug P/L | 0.0565 | 1.22 | 23.9 | TN | ug N/L | 0.658 | 8.45 | 453
2026-02-27 | 15:46 | OM_2-27-2026_12-05-56PM.OMN | S85 | 704 7/10/25 | 134 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.074 | 1.64 | 35.2 | TN | ug N/L | 0.675 | 8.67 | 467
2026-02-27 | 15:47 | OM_2-27-2026_12-05-56PM.OMN | DUP S85 |  | 140 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.0701 | 1.54 | 32.5 | TN | ug N/L | 0.697 | 8.97 | 486
2026-02-27 | 15:49 | OM_2-27-2026_12-05-56PM.OMN | S86 | 702 7/10/25 | 135 | Unknown | 1 | 1 | TP | ug P/L | 0.0659 | 1.68 | 36.4 | TN | ug N/L | 0.792 | 10.1 | 560
2026-02-27 | 15:50 | OM_2-27-2026_12-05-56PM.OMN | S87 | D2 7/31/25 | 136 | Unknown | 1 | 1 | TP | ug P/L | 0.739 | 13.4 | 468 | TN | ug N/L | 3.94 | 50.9 | 3380
2026-02-27 | 15:51 | OM_2-27-2026_12-05-56PM.OMN | S88 | A1 7/31/25 | 137 | Unknown | 1 | 1 | TP | ug P/L | 0.133 | 2.72 | 66 | TN | ug N/L | 0.628 | 8.11 | 432
2026-02-27 | 15:53 | OM_2-27-2026_12-05-56PM.OMN | S89 | C2 9/01/25 | 138 | Unknown | 1 | 1 | TP | ug P/L | 0.234 | 4.46 | 120 | TN | ug N/L | 2.4 | 30.4 | 1900
2026-02-27 | 15:54 | OM_2-27-2026_12-05-56PM.OMN | S90 | pinus virginiana 8/18/25 | 139 | Spike Set | 1 | 1 | TP | ug P/L | 0.104 | 2.13 | 49 | TN | ug N/L | 0.233 | 3.13 | 123
2026-02-27 | 15:55 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S90 |  | 141 | Spike Set | 1 | 1 | TP | ug P/L | 0.156 | 3.03 | 75.3 | TN | ug N/L | 0.538 | 6.96 | 360
2026-02-27 | 15:56 | OM_2-27-2026_12-05-56PM.OMN | BLANK |  | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.138 | 1.67 | 36.1 | TN | ug N/L | 0.0131 | 0.293 | -49.4
2026-02-27 | 15:58 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 |  | 18 | Check Standard | 1 | 1 | TP | ug P/L | 0.0513 | 1.1 | 20.5 | TN | ug N/L | 0.405 | 5.16 | 248
2026-02-27 | 15:59 | OM_2-27-2026_12-05-56PM.OMN | S91 | C1 9/29/25 | 142 | Unknown | 1 | 1 | TP | ug P/L | 0.585 | 9.06 | 283 | TN | ug N/L | 5.02 | 65.1 | 4470
2026-02-27 | 16:00 | OM_2-27-2026_12-05-56PM.OMN | S92 | R2 9/29/25 | 143 | Unknown | 1 | 1 | TP | ug P/L | 0.31 | 5.68 | 160 | TN | ug N/L | 0.422 | 5.44 | 266
2026-02-27 | 16:02 | OM_2-27-2026_12-05-56PM.OMN | S93 | A1 10/27/25 | 144 | Unknown | 1 | 1 | TP | ug P/L | 0.144 | 2.73 | 66.3 | TN | ug N/L | 0.94 | 12.2 | 688
2026-02-27 | 16:03 | OM_2-27-2026_12-05-56PM.OMN | S94 | C2 10/27/25 | 145 | Unknown | 1 | 1 | TP | ug P/L | 0.267 | 5.02 | 138 | TN | ug N/L | 1.6 | 20.3 | 1220
2026-02-27 | 16:04 | OM_2-27-2026_12-05-56PM.OMN | S95 | 707 9/15/25 | 146 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.392 | 7.24 | 215 | TN | ug N/L | 1.48 | 18.8 | 1120
2026-02-27 | 16:05 | OM_2-27-2026_12-05-56PM.OMN | DUP S95 |  | 152 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.397 | 7.3 | 217 | TN | ug N/L | 1.56 | 19.8 | 1190
2026-02-27 | 16:07 | OM_2-27-2026_12-05-56PM.OMN | S96 | C2 8/21/25 | 147 | Unknown | 1 | 1 | TP | ug P/L | 0.25 | 4.43 | 119 | TN | ug N/L | 1.72 | 21.8 | 1320
2026-02-27 | 16:08 | OM_2-27-2026_12-05-56PM.OMN | S97 | D4 7/31/25 | 148 | Unknown | 1 | 1 | TP | ug P/L | 0.215 | 4.16 | 110 | TN | ug N/L | 1.44 | 18.3 | 1080
2026-02-27 | 16:09 | OM_2-27-2026_12-05-56PM.OMN | S98 | A2 7/31/25 | 149 | Unknown | 1 | 1 | TP | ug P/L | 0.115 | 2.14 | 49.3 | TN | ug N/L | 1.39 | 17.6 | 1040
2026-02-27 | 16:11 | OM_2-27-2026_12-05-56PM.OMN | S99 | B4 8/21/25 | 150 | Unknown | 1 | 1 | TP | ug P/L | 0.0873 | 1.8 | 39.7 | TN | ug N/L | 0.724 | 9.29 | 506
2026-02-27 | 16:12 | OM_2-27-2026_12-05-56PM.OMN | S100 | B1 9/01/25 | 151 | Spike Set | 1 | 1 | TP | ug P/L | 0.0744 | 1.63 | 35 | TN | ug N/L | 0.77 | 9.93 | 547
2026-02-27 | 16:13 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S100 |  | 153 | Spike Set | 1 | 1 | TP | ug P/L | 0.123 | 2.54 | 60.9 | TN | ug N/L | 1.03 | 13.2 | 756
2026-02-27 | 16:14 | OM_2-27-2026_12-05-56PM.OMN | BISULFATE |  | S4 | Unknown | 1 | 1 | TP | ug P/L | 0.00999 | 0.107 | -5.29 | TN | ug N/L | 0.0191 | 0.317 | -48
2026-02-27 | 16:16 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 |  | 15 | Check Standard | 1 | 1 | TP | ug P/L | 0.0503 | 1.17 | 22.4 | TN | ug N/L | 0.384 | 5.05 | 241
2026-02-27 | 16:17 | OM_2-27-2026_12-05-56PM.OMN | NO2 |  | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.139 | 1.73 | 37.9 | TN | ug N/L | 1.86 | 24.4 | 1490
2026-02-27 | 16:18 | OM_2-27-2026_12-05-56PM.OMN | NO3 |  | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.147 | 1.87 | 41.7 | TN | ug N/L | 2.17 | 27.6 | 1710
2026-02-27 | 16:20 | OM_2-27-2026_12-05-56PM.OMN | S101 | B4 7/03/25 | 154 | Unknown | 1 | 1 | TP | ug P/L | 0.0663 | 1.42 | 29.3 | TN | ug N/L | 0.516 | 6.71 | 344
2026-02-27 | 16:21 | OM_2-27-2026_12-05-56PM.OMN | S102 | R1 8/18/25 | 155 | Unknown | 1 | 1 | TP | ug P/L | 0.0914 | 1.94 | 43.6 | TN | ug N/L | 0.448 | 5.86 | 292
2026-02-27 | 16:22 | OM_2-27-2026_12-05-56PM.OMN | S103 | A1 9/01/25 | 156 | Unknown | 1 | 1 | TP | ug P/L | 0.0866 | 1.84 | 41 | TN | ug N/L | 0.645 | 8.33 | 446
2026-02-27 | 16:23 | OM_2-27-2026_12-05-56PM.OMN | S104 | D4 10/27/25 | 157 | Unknown | 1 | 1 | TP | ug P/L | 0.101 | 2.1 | 48.2 | TN | ug N/L | 1.33 | 16.8 | 990
2026-02-27 | 16:25 | OM_2-27-2026_12-05-56PM.OMN | S105 | B4 9/01/25 | 158 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.064 | 1.38 | 28.2 | TN | ug N/L | 0.68 | 8.86 | 479
2026-02-27 | 16:26 | OM_2-27-2026_12-05-56PM.OMN | DUP S105 |  | 164 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.0876 | 1.38 | 28.3 | TN | ug N/L | 0.79 | 8.81 | 476
2026-02-27 | 16:27 | OM_2-27-2026_12-05-56PM.OMN | S106 | A2 9/01/25 | 159 | Unknown | 1 | 1 | TP | ug P/L | 0.128 | 2.59 | 62.2 | TN | ug N/L | 0.827 | 9.63 | 527
2026-02-27 | 16:28 | OM_2-27-2026_12-05-56PM.OMN | S107 | 705 8/07/25 | 160 | Unknown | 1 | 1 | TP | ug P/L | 0.0624 | 1.39 | 28.6 | TN | ug N/L | 0.608 | 7.91 | 419
2026-02-27 | 16:30 | OM_2-27-2026_12-05-56PM.OMN | S108 | 702 9/15/25 | 161 | Unknown | 1 | 1 | TP | ug P/L | 0.0848 | 1.79 | 39.3 | TN | ug N/L | 0.981 | 12.6 | 715
2026-02-27 | 16:31 | OM_2-27-2026_12-05-56PM.OMN | S109 | 701 8/07/25 | 162 | Unknown | 1 | 1 | TP | ug P/L | 0.0556 | 1.29 | 25.8 | TN | ug N/L | 0.604 | 7.81 | 413
2026-02-27 | 16:32 | OM_2-27-2026_12-05-56PM.OMN | S110 | C1 8/21/25 | 163 | Spike Set | 1 | 1 | TP | ug P/L | 0.242 | 4.63 | 125 | TN | ug N/L | 2.55 | 32.5 | 2050
2026-02-27 | 16:34 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S110 |  | 165 | Spike Set | 1 | 1 | TP | ug P/L | 0.18 | 3.51 | 89.8 | TN | ug N/L | 1.6 | 20.4 | 1230
2026-02-27 | 16:35 | OM_2-27-2026_12-05-56PM.OMN | BLANK |  | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.138 | 1.8 | 39.7 | TN | ug N/L | 0.00349 | 0.0185 | -66
2026-02-27 | 16:36 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 |  | 16 | Check Standard | 1 | 1 | TP | ug P/L | 0.0506 | 1.09 | 20.4 | TN | ug N/L | 0.354 | 4.71 | 221
2026-02-27 | 16:37 | OM_2-27-2026_12-05-56PM.OMN | S111 | R1 6/18/25 | 166 | Unknown | 1 | 1 | TP | ug P/L | 0.167 | 3.23 | 81.3 | TN | ug N/L | 0.389 | 5.07 | 243
2026-02-27 | 16:39 | OM_2-27-2026_12-05-56PM.OMN | S112 | D2 7/03/25 | 167 | Unknown | 1 | 1 | TP | ug P/L | 1.5 | 27.1 | 1250 | TN | ug N/L | 5.6 | 73.8 | 5170
2026-02-27 | 16:40 | OM_2-27-2026_12-05-56PM.OMN | S113 | C1 9/01/25 | 168 | Unknown | 1 | 1 | TP | ug P/L | 0.165 | 2.85 | 69.9 | TN | ug N/L | 1.39 | 17.8 | 1050
2026-02-27 | 16:41 | OM_2-27-2026_12-05-56PM.OMN | S114 | C2 7/31/25 | 169 | Unknown | 1 | 1 | TP | ug P/L | 0.0764 | 1.59 | 33.9 | TN | ug N/L | 0.744 | 7.83 | 415
2026-02-27 | 16:43 | OM_2-27-2026_12-05-56PM.OMN | S115 | C4 7/31/25 | 170 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.0814 | 1.76 | 38.7 | TN | ug N/L | 0.827 | 10.6 | 592
2026-02-27 | 16:44 | OM_2-27-2026_12-05-56PM.OMN | DUP S115 |  | 176 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.0568 | 1.41 | 29.1 | TN | ug N/L | 0.662 | 8.58 | 462
2026-02-27 | 16:45 | OM_2-27-2026_12-05-56PM.OMN | S116 | A3 8/21/25 | 171 | Unknown | 1 | 1 | TP | ug P/L | 0.12 | 2.44 | 58 | TN | ug N/L | 1.12 | 14.3 | 828
2026-02-27 | 16:46 | OM_2-27-2026_12-05-56PM.OMN | S117 | D1 7/31/25 | 172 | Unknown | 1 | 1 | TP | ug P/L | 0.862 | 15.7 | 576 | TN | ug N/L | 3.73 | 47.8 | 3150
2026-02-27 | 16:48 | OM_2-27-2026_12-05-56PM.OMN | S118 | D3 8/21/25 | 173 | Unknown | 1 | 1 | TP | ug P/L | 0.414 | 7.75 | 234 | TN | ug N/L | 2.18 | 27.8 | 1720
2026-02-27 | 16:49 | OM_2-27-2026_12-05-56PM.OMN | S119 | D3 10/27/25 | 174 | Unknown | 1 | 1 | TP | ug P/L | 0.0861 | 1.81 | 40 | TN | ug N/L | 0.768 | 9.92 | 546
2026-02-27 | 16:50 | OM_2-27-2026_12-05-56PM.OMN | S120 | D1 9/01/25 | 175 | Spike Set | 1 | 1 | TP | ug P/L | 0.244 | 4.72 | 128 | TN | ug N/L | 1.56 | 19.9 | 1190
2026-02-27 | 16:52 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S120 |  | 177 | Spike Set | 1 | 1 | TP | ug P/L | 0.207 | 3.97 | 104 | TN | ug N/L | 1.35 | 17.2 | 1020
2026-02-27 | 16:53 | OM_2-27-2026_12-05-56PM.OMN | BISULFATE |  | S4 | Unknown | 1 | 1 | TP | ug P/L | 0.00929 | 0.103 | -5.39 | TN | ug N/L | 0.0182 | 0.364 | -45.1
2026-02-27 | 16:54 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 |  | 17 | Check Standard | 1 | 1 | TP | ug P/L | 0.0507 | 1.15 | 21.9 | TN | ug N/L | 0.368 | 4.85 | 229
2026-02-27 | 16:55 | OM_2-27-2026_12-05-56PM.OMN | NO2 |  | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.137 | 1.66 | 35.8 | TN | ug N/L | 1.74 | 22.9 | 1390
2026-02-27 | 16:57 | OM_2-27-2026_12-05-56PM.OMN | NO3 |  | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.137 | 1.68 | 36.4 | TN | ug N/L | 1.9 | 24.4 | 1490
2026-02-27 | 16:58 | OM_2-27-2026_12-05-56PM.OMN | S121 | A4 6/18/25 | 178 | Unknown | 1 | 1 | TP | ug P/L | 0.318 | 6.06 | 173 | TN | ug N/L | 0.991 | 12.8 | 732
2026-02-27 | 16:59 | OM_2-27-2026_12-05-56PM.OMN | S122 | D3 6/18/25 | 179 | Unknown | 1 | 1 | TP | ug P/L | 0.373 | 7.05 | 208 | TN | ug N/L | 1.36 | 15.5 | 905
2026-02-27 | 17:01 | OM_2-27-2026_12-05-56PM.OMN | S123 | D1 7/03/25 | 180 | Unknown | 1 | 1 | TP | ug P/L | 0.778 | 14.4 | 514 | TN | ug N/L | 3.09 | 39.2 | 2520
2026-02-27 | 17:02 | OM_2-27-2026_12-05-56PM.OMN | S124 | B1 6/18/25 | 181 | Unknown | 1 | 1 | TP | ug P/L | 0.543 | 10.1 | 325 | TN | ug N/L | 1.69 | 21.4 | 1290
2026-02-27 | 17:03 | OM_2-27-2026_12-05-56PM.OMN | S125 | C3 9/29/25 | 182 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.42 | 7.93 | 240 | TN | ug N/L | 3.21 | 40.8 | 2630
2026-02-27 | 17:04 | OM_2-27-2026_12-05-56PM.OMN | DUP S125 |  | 188 | Duplicate Set | 1 | 1 | TP | ug P/L | 0.254 | 4.89 | 133 | TN | ug N/L | 2.01 | 25.6 | 1570
2026-02-27 | 17:06 | OM_2-27-2026_12-05-56PM.OMN | S126 | D4 6/18/25 | 183 | Unknown | 1 | 1 | TP | ug P/L | 0.406 | 7.25 | 215 | TN | ug N/L | 1.19 | 15.2 | 885
2026-02-27 | 17:07 | OM_2-27-2026_12-05-56PM.OMN | S127 | A3 6/18/25 | 184 | Unknown | 1 | 1 | TP | ug P/L | 0.889 | 16.4 | 614 | TN | ug N/L | 4.3 | 55.1 | 3690
2026-02-27 | 17:08 | OM_2-27-2026_12-05-56PM.OMN | S128 | C3 6/18/25 | 185 | Unknown | 1 | 1 | TP | ug P/L | 0.316 | 5.99 | 170 | TN | ug N/L | 1.1 | 14 | 806
2026-02-27 | 17:10 | OM_2-27-2026_12-05-56PM.OMN | S129 | R1 9/29/25 | 186 | Unknown | 1 | 1 | TP | ug P/L | 0.72 | 13.3 | 464 | TN | ug N/L | 0.955 | 12.2 | 689
2026-02-27 | 17:11 | OM_2-27-2026_12-05-56PM.OMN | S130 | D2 6/18/25 | 187 | Spike Set | 1 | 1 | TP | ug P/L | 0.852 | 15.7 | 579 | TN | ug N/L | 4.66 | 60.3 | 4090
2026-02-27 | 17:12 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S130 |  | 189 | Spike Set | 1 | 1 | TP | ug P/L | 0.754 | 13.9 | 493 | TN | ug N/L | 3.89 | 50.1 | 3320
2026-02-27 | 17:14 | OM_2-27-2026_12-05-56PM.OMN | BLANK |  | S3 | Unknown | 1 | 1 | TP | ug P/L | 0.136 | 1.69 | 36.7 | TN | ug N/L | 0.00781 | 0.181 | -56.2
2026-02-27 | 17:15 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 |  | 18 | Check Standard | 1 | 1 | TP | ug P/L | 0.0504 | 1.13 | 21.4 | TN | ug N/L | 0.339 | 4.47 | 206
2026-02-27 | 17:16 | OM_2-27-2026_12-05-56PM.OMN | S131 | B2 6/18/25 | 190 | Unknown | 1 | 1 | TP | ug P/L | 0.854 | 15.8 | 584 | TN | ug N/L | 2.99 | 38.1 | 2440
2026-02-27 | 17:17 | OM_2-27-2026_12-05-56PM.OMN | S132 | A1 7/03/25 | 191 | Unknown | 1 | 1 | TP | ug P/L | 0.143 | 2.91 | 71.6 | TN | ug N/L | 0.666 | 8.6 | 463
2026-02-27 | 17:19 | OM_2-27-2026_12-05-56PM.OMN | S133 | B4 9/29/25 | 192 | Unknown | 1 | 1 | TP | ug P/L | 0.142 | 2.85 | 69.9 | TN | ug N/L | 1.11 | 14.2 | 819
2026-02-27 | 17:20 | OM_2-27-2026_12-05-56PM.OMN | S134 | D1 6/18/25 | 193 | Unknown | 1 | 1 | TP | ug P/L | 0.558 | 10.5 | 342 | TN | ug N/L | 1.56 | 19.8 | 1180
2026-02-27 | 17:21 | OM_2-27-2026_12-05-56PM.OMN | S135 | A2 6/18/25 | 194 | Duplicate Set | 1 | 1 | TP | ug P/L | 1.13 | 20.5 | 834 | TN | ug N/L | 2.58 | 32.5 | 2050
2026-02-27 | 17:23 | OM_2-27-2026_12-05-56PM.OMN | DUP S135 |  | 199 | Duplicate Set | 1 | 1 | TP | ug P/L | 1.4 | 25.5 | 1140 | TN | ug N/L | 3.43 | 43.4 | 2820
2026-02-27 | 17:24 | OM_2-27-2026_12-05-56PM.OMN | S136 | R2 10/27/25 | 195 | Unknown | 1 | 1 | TP | ug P/L | 0.232 | 4.68 | 127 | TN | ug N/L | 0.316 | 4.1 | 183
2026-02-27 | 17:25 | OM_2-27-2026_12-05-56PM.OMN | S137 | A1 9/29/25 | 196 | Unknown | 1 | 1 | TP | ug P/L | 0.0791 | 1.61 | 34.6 | TN | ug N/L | 0.642 | 8.29 | 443
2026-02-27 | 17:26 | OM_2-27-2026_12-05-56PM.OMN | S138 | B1 7/03/25 | 197 | Unknown | 1 | 1 | TP | ug P/L | 0.134 | 2.66 | 64.2 | TN | ug N/L | 1.09 | 13.9 | 802
2026-02-27 | 17:28 | OM_2-27-2026_12-05-56PM.OMN | S139 | C4 6/18/25 | 198 | Spike Set | 1 | 1 | TP | ug P/L | 0.382 | 7.05 | 208 | TN | ug N/L | 1.17 | 15.1 | 874
2026-02-27 | 17:29 | OM_2-27-2026_12-05-56PM.OMN | SPIKE S139 |  | 200 | Spike Set | 1 | 1 | TP | ug P/L | 0.443 | 8.08 | 246 | TN | ug N/L | 1.49 | 19.1 | 1140
2026-02-27 | 17:30 | OM_2-27-2026_12-05-56PM.OMN | CHECK 150 |  | 1 | Check Standard | 1 | 1 | TP | ug P/L | 0.3 | 5.46 | 152 | TN | ug N/L | 2.4 | 21.4 | 1290
2026-02-27 | 17:32 | OM_2-27-2026_12-05-56PM.OMN | CHECK 100 |  | 3 | Check Standard | 1 | 1 | TP | ug P/L | 0.197 | 3.69 | 95.3 | TN | ug N/L | 1.27 | 16.2 | 946
2026-02-27 | 17:33 | OM_2-27-2026_12-05-56PM.OMN | CHECK 50 |  | 5 | Check Standard | 1 | 1 | TP | ug P/L | 0.124 | 2.05 | 46.8 | TN | ug N/L | 0.652 | 8.4 | 450
2026-02-27 | 17:34 | OM_2-27-2026_12-05-56PM.OMN | CHECK 25 |  | 7 | Check Standard | 1 | 1 | TP | ug P/L | 0.0529 | 1.19 | 23.1 | TN | ug N/L | 0.397 | 5.19 | 250
2026-02-27 | 17:35 | OM_2-27-2026_12-05-56PM.OMN | CHECK 0 |  | 13 | Check Standard | 1 | 1 | TP | ug P/L | 0.00922 | 0.152 | -4.16 | TN | ug N/L | 0.0692 | 1.07 | -2.44
2026-02-27 | 17:37 | OM_2-27-2026_12-05-56PM.OMN | NO2 |  | S1 | Unknown | 1 | 1 | TP | ug P/L | 0.137 | 1.69 | 36.6 | TN | ug N/L | 1.83 | 23.9 | 1460
2026-02-27 | 17:40 | OM_2-27-2026_12-05-56PM.OMN | NO3 |  | S2 | Unknown | 1 | 1 | TP | ug P/L | 0.135 | 1.66 | 36 | TN | ug N/L | 1.86 | 23.7 | 1450